## calibration_GML_2026_001
etalon_temperature | etalon_pressure | etalon_humidity | logger_temperature | logger_pressure | logger_humidity | calibration_delta_temperature | calibration_delta_pressure | calibration_delta_humidity | model_temperature | model_pressure | model_humidity | delta_model_temperature | delta_model_pressure | delta_model_humidity
35.93 | 996.65 | 29.54 | 35.95 | 998.41 | 38.75 | -0.02 | -1.76 | -9.21 | 35.6094 | 996.6635 | 30.7728 | 0.32064 | -0.013476 | -1.2328
35.96 | 996.68 | 29.65 | 36.02 | 998.45 | 38.71 | -0.06 | -1.77 | -9.06 | 35.6926 | 996.7035 | 30.7328 | 0.26739 | -0.023476 | -1.0828
36 | 996.66 | 29.68 | 36.12 | 998.45 | 38.56 | -0.12 | -1.79 | -8.88 | 35.8115 | 996.7035 | 30.5828 | 0.18847 | -0.043476 | -0.90281
36.05 | 996.71 | 29.63 | 36.2 | 998.49 | 38.36 | -0.15 | -1.78 | -8.73 | 35.9067 | 996.7435 | 30.3828 | 0.14333 | -0.033476 | -0.75281
36.09 | 996.69 | 29.6 | 36.22 | 998.47 | 38.28 | -0.13 | -1.78 | -8.68 | 35.9305 | 996.7235 | 30.3028 | 0.15954 | -0.033476 | -0.70281
36.15 | 996.72 | 29.39 | 36.27 | 998.5 | 38.18 | -0.12 | -1.78 | -8.79 | 35.9899 | 996.7535 | 30.2028 | 0.16008 | -0.033476 | -0.81281
36.19 | 996.74 | 29.29 | 36.47 | 998.53 | 37.95 | -0.28 | -1.79 | -8.66 | 36.2278 | 996.7835 | 29.9728 | -0.03777 | -0.043476 | -0.68281
36.33 | 996.74 | 29.07 | 36.65 | 998.54 | 37.58 | -0.32 | -1.8 | -8.51 | 36.4418 | 996.7935 | 29.6028 | -0.11184 | -0.053476 | -0.53281
36.48 | 996.75 | 28.8 | 36.79 | 998.54 | 37.36 | -0.31 | -1.79 | -8.56 | 36.6083 | 996.7935 | 29.3828 | -0.12833 | -0.043476 | -0.58281
36.61 | 996.76 | 28.59 | 36.89 | 998.53 | 37.16 | -0.28 | -1.77 | -8.57 | 36.7273 | 996.7835 | 29.1828 | -0.11726 | -0.023476 | -0.59281
36.68 | 996.78 | 28.42 | 36.93 | 998.56 | 37.06 | -0.25 | -1.78 | -8.64 | 36.7748 | 996.8135 | 29.0828 | -0.094827 | -0.033476 | -0.66281
36.77 | 996.76 | 28.29 | 37.04 | 998.55 | 36.9 | -0.27 | -1.79 | -8.61 | 36.9056 | 996.8035 | 28.9228 | -0.13564 | -0.043476 | -0.63281
36.85 | 996.76 | 28.14 | 37.08 | 998.53 | 36.83 | -0.23 | -1.77 | -8.69 | 36.9532 | 996.7835 | 28.8528 | -0.10322 | -0.023476 | -0.71281
36.92 | 996.76 | 28.05 | 37.17 | 998.52 | 36.69 | -0.25 | -1.76 | -8.64 | 37.0602 | 996.7735 | 28.7128 | -0.14025 | -0.013476 | -0.66281
36.97 | 996.75 | 27.94 | 37.23 | 998.52 | 36.64 | -0.26 | -1.77 | -8.7 | 37.1316 | 996.7735 | 28.6628 | -0.1616 | -0.023476 | -0.72281
37 | 996.76 | 27.86 | 37.22 | 998.53 | 36.58 | -0.22 | -1.77 | -8.72 | 37.1197 | 996.7835 | 28.6028 | -0.11971 | -0.023476 | -0.74281
37.04 | 996.77 | 27.78 | 37.26 | 998.52 | 36.51 | -0.22 | -1.75 | -8.73 | 37.1673 | 996.7735 | 28.5328 | -0.12728 | -0.0034762 | -0.75281
37.06 | 996.77 | 27.74 | 37.28 | 998.51 | 36.47 | -0.22 | -1.74 | -8.73 | 37.1911 | 996.7635 | 28.4928 | -0.13107 | 0.0065238 | -0.75281
37.1 | 996.77 | 27.66 | 37.34 | 998.5 | 36.4 | -0.24 | -1.73 | -8.74 | 37.2624 | 996.7535 | 28.4228 | -0.16242 | 0.016524 | -0.76281
37.13 | 996.75 | 27.6 | 37.32 | 998.47 | 36.38 | -0.19 | -1.72 | -8.78 | 37.2386 | 996.7235 | 28.4028 | -0.10864 | 0.026524 | -0.80281
37.17 | 996.75 | 27.57 | 37.37 | 998.49 | 36.32 | -0.2 | -1.74 | -8.75 | 37.2981 | 996.7435 | 28.3428 | -0.1281 | 0.0065238 | -0.77281
37.19 | 996.76 | 27.51 | 37.42 | 998.49 | 36.22 | -0.23 | -1.73 | -8.71 | 37.3576 | 996.7435 | 28.2428 | -0.16756 | 0.016524 | -0.73281
37.23 | 996.76 | 27.44 | 37.42 | 998.5 | 36.23 | -0.19 | -1.74 | -8.79 | 37.3576 | 996.7535 | 28.2528 | -0.12756 | 0.0065238 | -0.81281
37.24 | 996.76 | 27.41 | 37.45 | 998.5 | 36.2 | -0.21 | -1.74 | -8.79 | 37.3932 | 996.7535 | 28.2228 | -0.15324 | 0.0065238 | -0.81281
37.27 | 996.72 | 27.37 | 37.48 | 998.47 | 36.15 | -0.21 | -1.75 | -8.78 | 37.4289 | 996.7235 | 28.1728 | -0.15892 | -0.0034762 | -0.80281
37.32 | 996.72 | 27.31 | 37.47 | 998.47 | 36.11 | -0.15 | -1.75 | -8.8 | 37.417 | 996.7235 | 28.1328 | -0.097024 | -0.0034762 | -0.82281
37.33 | 996.72 | 27.27 | 37.51 | 998.46 | 36.04 | -0.18 | -1.74 | -8.77 | 37.4646 | 996.7135 | 28.0628 | -0.13459 | 0.0065238 | -0.79281
37.34 | 996.73 | 27.25 | 37.52 | 998.47 | 36.04 | -0.18 | -1.74 | -8.79 | 37.4765 | 996.7235 | 28.0628 | -0.13649 | 0.0065238 | -0.81281
37.36 | 996.73 | 27.22 | 37.57 | 998.48 | 35.98 | -0.21 | -1.75 | -8.76 | 37.5359 | 996.7335 | 28.0028 | -0.17595 | -0.0034762 | -0.78281
37.4 | 996.73 | 27.15 | 37.57 | 998.45 | 35.95 | -0.17 | -1.72 | -8.8 | 37.5359 | 996.7035 | 27.9728 | -0.13595 | 0.026524 | -0.82281
37.4 | 996.75 | 27.13 | 37.56 | 998.5 | 35.97 | -0.16 | -1.75 | -8.84 | 37.5241 | 996.7535 | 27.9928 | -0.12406 | -0.0034762 | -0.86281
37.43 | 996.74 | 27.11 | 37.6 | 998.47 | 35.88 | -0.17 | -1.73 | -8.77 | 37.5716 | 996.7235 | 27.9028 | -0.14163 | 0.016524 | -0.79281
37.44 | 996.75 | 27.06 | 37.61 | 998.49 | 35.91 | -0.17 | -1.74 | -8.85 | 37.5835 | 996.7435 | 27.9328 | -0.14352 | 0.0065238 | -0.87281
37.49 | 996.76 | 27.02 | 37.66 | 998.49 | 35.81 | -0.17 | -1.73 | -8.79 | 37.643 | 996.7435 | 27.8328 | -0.15298 | 0.016524 | -0.81281
37.53 | 996.75 | 26.97 | 37.68 | 998.48 | 35.8 | -0.15 | -1.73 | -8.83 | 37.6668 | 996.7335 | 27.8228 | -0.13677 | 0.016524 | -0.85281
37.53 | 996.75 | 26.92 | 37.66 | 998.49 | 35.79 | -0.13 | -1.74 | -8.87 | 37.643 | 996.7435 | 27.8128 | -0.11298 | 0.0065238 | -0.89281
37.5 | 996.77 | 26.95 | 37.68 | 998.49 | 35.77 | -0.18 | -1.72 | -8.82 | 37.6668 | 996.7435 | 27.7928 | -0.16677 | 0.026524 | -0.84281
37.52 | 996.76 | 26.93 | 37.69 | 998.5 | 35.77 | -0.17 | -1.74 | -8.84 | 37.6787 | 996.7535 | 27.7928 | -0.15866 | 0.0065238 | -0.86281
37.55 | 996.75 | 26.9 | 37.72 | 998.51 | 35.69 | -0.17 | -1.76 | -8.79 | 37.7143 | 996.7635 | 27.7128 | -0.16434 | -0.013476 | -0.81281
37.59 | 996.76 | 26.86 | 37.76 | 998.5 | 35.65 | -0.17 | -1.74 | -8.79 | 37.7619 | 996.7535 | 27.6728 | -0.17191 | 0.0065238 | -0.81281
37.6 | 996.74 | 26.82 | 37.75 | 998.48 | 35.62 | -0.15 | -1.74 | -8.8 | 37.75 | 996.7335 | 27.6428 | -0.15002 | 0.0065238 | -0.82281
37.61 | 996.74 | 26.8 | 37.76 | 998.48 | 35.64 | -0.15 | -1.74 | -8.84 | 37.7619 | 996.7335 | 27.6628 | -0.15191 | 0.0065238 | -0.86281
37.61 | 996.74 | 26.79 | 37.79 | 998.46 | 35.64 | -0.18 | -1.72 | -8.85 | 37.7976 | 996.7135 | 27.6628 | -0.18759 | 0.026524 | -0.87281
37.65 | 996.73 | 26.74 | 37.81 | 998.48 | 35.53 | -0.16 | -1.75 | -8.79 | 37.8214 | 996.7335 | 27.5528 | -0.17137 | -0.0034762 | -0.81281
37.68 | 996.72 | 26.71 | 37.8 | 998.49 | 35.56 | -0.12 | -1.77 | -8.85 | 37.8095 | 996.7435 | 27.5828 | -0.12948 | -0.023476 | -0.87281
37.69 | 996.72 | 26.69 | 37.82 | 998.48 | 35.48 | -0.13 | -1.76 | -8.79 | 37.8333 | 996.7335 | 27.5028 | -0.14326 | -0.013476 | -0.81281
37.72 | 996.71 | 26.66 | 37.85 | 998.46 | 35.45 | -0.13 | -1.75 | -8.79 | 37.8689 | 996.7135 | 27.4728 | -0.14894 | -0.0034762 | -0.81281
37.74 | 996.7 | 26.63 | 37.88 | 998.46 | 35.47 | -0.14 | -1.76 | -8.84 | 37.9046 | 996.7135 | 27.4928 | -0.16462 | -0.013476 | -0.86281
37.71 | 996.69 | 26.66 | 37.82 | 998.43 | 35.95 | -0.11 | -1.74 | -9.29 | 37.8333 | 996.6835 | 27.9728 | -0.12326 | 0.0065238 | -1.3128
37.6 | 996.69 | 27.03 | 37.67 | 998.43 | 36 | -0.07 | -1.74 | -8.97 | 37.6549 | 996.6835 | 28.0228 | -0.054875 | 0.0065238 | -0.99281
37.39 | 996.68 | 26.93 | 37.51 | 998.42 | 35.94 | -0.12 | -1.74 | -9.01 | 37.4646 | 996.6735 | 27.9628 | -0.074594 | 0.0065238 | -1.0328
36.93 | 996.68 | 27.04 | 37.18 | 998.43 | 35.87 | -0.25 | -1.75 | -8.83 | 37.0721 | 996.6835 | 27.8928 | -0.14214 | -0.0034762 | -0.85281
36.24 | 996.68 | 27.07 | 36.4 | 998.47 | 36.61 | -0.16 | -1.79 | -9.54 | 36.1445 | 996.7235 | 28.6328 | 0.095478 | -0.043476 | -1.5628
35.38 | 996.68 | 27.92 | 35.98 | 998.46 | 36.66 | -0.6 | -1.78 | -8.74 | 35.645 | 996.7135 | 28.6828 | -0.26504 | -0.033476 | -0.76281
34.68 | 996.67 | 28.7 | 35.53 | 998.44 | 37.27 | -0.85 | -1.77 | -8.57 | 35.1099 | 996.6935 | 29.2928 | -0.42987 | -0.023476 | -0.59281
34.19 | 996.67 | 29.1 | 35.18 | 998.45 | 37.33 | -0.99 | -1.78 | -8.23 | 34.6936 | 996.7035 | 29.3528 | -0.50363 | -0.033476 | -0.25281
33.71 | 996.66 | 29.48 | 34.76 | 998.42 | 37.5 | -1.05 | -1.76 | -8.02 | 34.1941 | 996.6735 | 29.5228 | -0.48415 | -0.013476 | -0.04281
33.35 | 996.66 | 30.06 | 34.73 | 998.44 | 37.65 | -1.38 | -1.78 | -7.59 | 34.1585 | 996.6935 | 29.6728 | -0.80847 | -0.033476 | 0.38719
33.21 | 996.67 | 30.14 | 34.51 | 998.43 | 37.63 | -1.3 | -1.76 | -7.49 | 33.8968 | 996.6835 | 29.6528 | -0.68683 | -0.013476 | 0.48719
32.91 | 996.69 | 30.4 | 34.28 | 998.43 | 37.8 | -1.37 | -1.74 | -7.4 | 33.6233 | 996.6835 | 29.8228 | -0.7133 | 0.0065238 | 0.57719
32.78 | 996.67 | 30.74 | 34.09 | 998.42 | 38.04 | -1.31 | -1.75 | -7.3 | 33.3973 | 996.6735 | 30.0628 | -0.61735 | -0.0034762 | 0.67719
32.74 | 996.66 | 30.66 | 34.03 | 998.42 | 38.07 | -1.29 | -1.76 | -7.41 | 33.326 | 996.6735 | 30.0928 | -0.58599 | -0.013476 | 0.56719
32.57 | 996.67 | 30.7 | 33.86 | 998.44 | 38.02 | -1.29 | -1.77 | -7.32 | 33.1238 | 996.6935 | 30.0428 | -0.55382 | -0.023476 | 0.65719
32.33 | 996.67 | 30.9 | 33.57 | 998.41 | 38.11 | -1.24 | -1.74 | -7.21 | 32.7789 | 996.6635 | 30.1328 | -0.44893 | 0.0065238 | 0.76719
32.29 | 996.66 | 30.88 | 33.59 | 998.41 | 37.92 | -1.3 | -1.75 | -7.04 | 32.8027 | 996.6635 | 29.9428 | -0.51272 | -0.0034762 | 0.93719
32.17 | 996.65 | 30.91 | 33.31 | 998.39 | 38.25 | -1.14 | -1.74 | -7.34 | 32.4697 | 996.6435 | 30.2728 | -0.29973 | 0.0065238 | 0.63719
32.02 | 996.66 | 31 | 33.22 | 998.4 | 38.15 | -1.2 | -1.74 | -7.15 | 32.3627 | 996.6535 | 30.1728 | -0.3427 | 0.0065238 | 0.82719
31.97 | 996.66 | 31.02 | 33.08 | 998.41 | 38.05 | -1.11 | -1.75 | -7.03 | 32.1962 | 996.6635 | 30.0728 | -0.2262 | -0.0034762 | 0.94719
31.86 | 996.64 | 30.96 | 33.09 | 998.37 | 38.19 | -1.23 | -1.73 | -7.23 | 32.2081 | 996.6235 | 30.2128 | -0.34809 | 0.016524 | 0.74719
31.85 | 996.64 | 30.82 | 32.99 | 998.37 | 38.05 | -1.14 | -1.73 | -7.23 | 32.0892 | 996.6235 | 30.0728 | -0.23917 | 0.016524 | 0.74719
31.66 | 996.66 | 30.93 | 32.78 | 998.39 | 38.31 | -1.12 | -1.73 | -7.38 | 31.8394 | 996.6435 | 30.3328 | -0.17942 | 0.016524 | 0.59719
31.62 | 996.64 | 30.84 | 32.72 | 998.39 | 38.11 | -1.1 | -1.75 | -7.27 | 31.7681 | 996.6435 | 30.1328 | -0.14807 | -0.0034762 | 0.70719
31.47 | 996.64 | 30.99 | 32.72 | 998.39 | 38.23 | -1.25 | -1.75 | -7.24 | 31.7681 | 996.6435 | 30.2528 | -0.29807 | -0.0034762 | 0.73719
31.47 | 996.63 | 30.96 | 32.66 | 998.37 | 38.15 | -1.19 | -1.74 | -7.19 | 31.6967 | 996.6235 | 30.1728 | -0.22671 | 0.0065238 | 0.78719
31.5 | 996.62 | 30.79 | 32.55 | 998.38 | 37.94 | -1.05 | -1.76 | -7.15 | 31.5659 | 996.6335 | 29.9628 | -0.065895 | -0.013476 | 0.82719
31.39 | 996.61 | 30.75 | 32.42 | 998.36 | 38.05 | -1.03 | -1.75 | -7.3 | 31.4113 | 996.6135 | 30.0728 | -0.021292 | -0.0034762 | 0.67719
31.2 | 996.61 | 30.85 | 32.4 | 998.37 | 38.32 | -1.2 | -1.76 | -7.47 | 31.3875 | 996.6235 | 30.3428 | -0.18751 | -0.013476 | 0.50719
31.18 | 996.6 | 30.93 | 32.41 | 998.35 | 37.95 | -1.23 | -1.75 | -7.02 | 31.3994 | 996.6035 | 29.9728 | -0.2194 | -0.0034762 | 0.95719
31.24 | 996.6 | 31.26 | 32.47 | 998.37 | 38.18 | -1.23 | -1.77 | -6.92 | 31.4708 | 996.6235 | 30.2028 | -0.23075 | -0.023476 | 1.0572
31.31 | 996.59 | 30.89 | 32.43 | 998.35 | 38.13 | -1.12 | -1.76 | -7.24 | 31.4232 | 996.6035 | 30.1528 | -0.11318 | -0.013476 | 0.73719
31.27 | 996.62 | 30.73 | 32.3 | 998.38 | 38.13 | -1.03 | -1.76 | -7.4 | 31.2686 | 996.6335 | 30.1528 | 0.0014186 | -0.013476 | 0.57719
31.18 | 996.62 | 30.84 | 32.33 | 998.38 | 37.93 | -1.15 | -1.76 | -7.09 | 31.3043 | 996.6335 | 29.9528 | -0.12426 | -0.013476 | 0.88719
31.12 | 996.62 | 30.75 | 32.21 | 998.4 | 37.96 | -1.09 | -1.78 | -7.21 | 31.1615 | 996.6535 | 29.9828 | -0.041549 | -0.033476 | 0.76719
31.01 | 996.62 | 30.77 | 32.1 | 998.4 | 38.01 | -1.09 | -1.78 | -7.24 | 31.0307 | 996.6535 | 30.0328 | -0.020731 | -0.033476 | 0.73719
30.95 | 996.61 | 30.9 | 32.06 | 998.39 | 38.11 | -1.11 | -1.78 | -7.21 | 30.9832 | 996.6435 | 30.1328 | -0.033161 | -0.033476 | 0.76719
30.92 | 996.6 | 30.86 | 32.08 | 998.37 | 38.04 | -1.16 | -1.77 | -7.18 | 31.0069 | 996.6235 | 30.0628 | -0.086946 | -0.023476 | 0.79719
30.99 | 996.58 | 30.86 | 32.15 | 998.37 | 37.91 | -1.16 | -1.79 | -7.05 | 31.0902 | 996.6235 | 29.9328 | -0.10019 | -0.043476 | 0.92719
31.09 | 996.57 | 30.93 | 32.3 | 998.33 | 37.7 | -1.21 | -1.76 | -6.77 | 31.2686 | 996.5835 | 29.7228 | -0.17858 | -0.013476 | 1.2072
31.1 | 996.57 | 30.76 | 32.23 | 998.35 | 37.85 | -1.13 | -1.78 | -7.09 | 31.1853 | 996.6035 | 29.8728 | -0.085334 | -0.033476 | 0.88719
31.06 | 996.58 | 30.68 | 32.14 | 998.36 | 37.85 | -1.08 | -1.78 | -7.17 | 31.0783 | 996.6135 | 29.8728 | -0.018301 | -0.033476 | 0.80719
30.98 | 996.58 | 30.65 | 32.01 | 998.36 | 37.94 | -1.03 | -1.78 | -7.29 | 30.9237 | 996.6135 | 29.9628 | 0.056302 | -0.033476 | 0.68719
30.95 | 996.57 | 30.63 | 32.11 | 998.34 | 37.72 | -1.16 | -1.77 | -7.09 | 31.0426 | 996.5935 | 29.7428 | -0.092623 | -0.023476 | 0.88719
30.91 | 996.57 | 30.67 | 31.94 | 998.34 | 37.91 | -1.03 | -1.77 | -7.24 | 30.8405 | 996.5935 | 29.9328 | 0.06955 | -0.023476 | 0.73719
30.86 | 996.56 | 30.68 | 31.91 | 998.33 | 37.96 | -1.05 | -1.77 | -7.28 | 30.8048 | 996.5835 | 29.9828 | 0.055228 | -0.023476 | 0.69719
30.87 | 996.55 | 30.66 | 31.91 | 998.32 | 37.91 | -1.04 | -1.77 | -7.25 | 30.8048 | 996.5735 | 29.9328 | 0.065228 | -0.023476 | 0.72719
30.87 | 996.55 | 30.61 | 31.88 | 998.32 | 37.85 | -1.01 | -1.77 | -7.24 | 30.7691 | 996.5735 | 29.8728 | 0.10091 | -0.023476 | 0.73719
30.83 | 996.53 | 30.66 | 31.92 | 998.29 | 37.88 | -1.09 | -1.76 | -7.22 | 30.8167 | 996.5435 | 29.9028 | 0.013335 | -0.013476 | 0.75719
30.79 | 996.53 | 30.62 | 31.83 | 998.29 | 37.88 | -1.04 | -1.76 | -7.26 | 30.7096 | 996.5435 | 29.9028 | 0.080368 | -0.013476 | 0.71719
30.8 | 996.5 | 30.6 | 31.83 | 998.26 | 37.79 | -1.03 | -1.76 | -7.19 | 30.7096 | 996.5135 | 29.8128 | 0.090368 | -0.013476 | 0.78719
30.79 | 996.49 | 30.61 | 31.83 | 998.25 | 37.82 | -1.04 | -1.76 | -7.21 | 30.7096 | 996.5035 | 29.8428 | 0.080368 | -0.013476 | 0.76719
30.76 | 996.48 | 30.52 | 31.66 | 998.26 | 37.94 | -0.9 | -1.78 | -7.42 | 30.5075 | 996.5135 | 29.9628 | 0.25254 | -0.033476 | 0.55719
30.65 | 996.51 | 30.7 | 31.61 | 998.27 | 37.93 | -0.96 | -1.76 | -7.23 | 30.448 | 996.5235 | 29.9528 | 0.202 | -0.013476 | 0.74719
30.6 | 996.5 | 30.73 | 31.62 | 998.28 | 38.05 | -1.02 | -1.78 | -7.32 | 30.4599 | 996.5335 | 30.0728 | 0.14011 | -0.033476 | 0.65719
30.54 | 996.5 | 30.82 | 31.57 | 998.28 | 38.01 | -1.03 | -1.78 | -7.19 | 30.4004 | 996.5335 | 30.0328 | 0.13957 | -0.033476 | 0.78719
30.45 | 996.49 | 30.91 | 31.46 | 998.26 | 38.19 | -1.01 | -1.77 | -7.28 | 30.2696 | 996.5135 | 30.2128 | 0.18039 | -0.023476 | 0.69719
30.39 | 996.5 | 30.98 | 31.42 | 998.28 | 38.12 | -1.03 | -1.78 | -7.14 | 30.222 | 996.5335 | 30.1428 | 0.16796 | -0.033476 | 0.83719
30.44 | 996.51 | 30.93 | 31.46 | 998.28 | 38.15 | -1.02 | -1.77 | -7.22 | 30.2696 | 996.5335 | 30.1728 | 0.17039 | -0.023476 | 0.75719
30.36 | 996.53 | 31.02 | 31.39 | 998.31 | 38.3 | -1.03 | -1.78 | -7.28 | 30.1864 | 996.5635 | 30.3228 | 0.17364 | -0.033476 | 0.69719
30.39 | 996.53 | 31.06 | 31.44 | 998.3 | 38.22 | -1.05 | -1.77 | -7.16 | 30.2458 | 996.5535 | 30.2428 | 0.14418 | -0.023476 | 0.81719
30.48 | 996.53 | 30.87 | 31.44 | 998.32 | 38.18 | -0.96 | -1.79 | -7.31 | 30.2458 | 996.5735 | 30.2028 | 0.23418 | -0.043476 | 0.66719
30.57 | 996.54 | 31.63 | 31.69 | 998.33 | 38.38 | -1.12 | -1.79 | -6.75 | 30.5431 | 996.5835 | 30.4028 | 0.026863 | -0.043476 | 1.2272
30.97 | 996.55 | 30.68 | 31.94 | 998.31 | 37.89 | -0.97 | -1.76 | -7.21 | 30.8405 | 996.5635 | 29.9128 | 0.12955 | -0.013476 | 0.76719
31.47 | 996.56 | 30.21 | 32.25 | 998.32 | 37.46 | -0.78 | -1.76 | -7.25 | 31.2091 | 996.5735 | 29.4828 | 0.26088 | -0.013476 | 0.72719
31.93 | 996.55 | 29.56 | 32.61 | 998.31 | 36.92 | -0.68 | -1.76 | -7.36 | 31.6373 | 996.5635 | 28.9428 | 0.29275 | -0.013476 | 0.61719
32.35 | 996.57 | 28.92 | 32.93 | 998.3 | 36.47 | -0.58 | -1.73 | -7.55 | 32.0178 | 996.5535 | 28.4928 | 0.33219 | 0.016524 | 0.42719
32.75 | 996.55 | 28.41 | 33.24 | 998.27 | 36.09 | -0.49 | -1.72 | -7.68 | 32.3865 | 996.5235 | 28.1128 | 0.36352 | 0.026524 | 0.29719
33.11 | 996.53 | 27.95 | 33.52 | 998.26 | 35.77 | -0.41 | -1.73 | -7.82 | 32.7195 | 996.5135 | 27.7928 | 0.39053 | 0.016524 | 0.15719
33.44 | 996.53 | 27.46 | 33.8 | 998.24 | 35.45 | -0.36 | -1.71 | -7.99 | 33.0525 | 996.4935 | 27.4728 | 0.38754 | 0.036524 | -0.01281
33.72 | 996.52 | 27.1 | 34.03 | 998.25 | 35.17 | -0.31 | -1.73 | -8.07 | 33.326 | 996.5035 | 27.1928 | 0.39401 | 0.016524 | -0.09281
34 | 996.51 | 26.75 | 34.25 | 998.24 | 34.9 | -0.25 | -1.73 | -8.15 | 33.5876 | 996.4935 | 26.9228 | 0.41237 | 0.016524 | -0.17281
34.22 | 996.53 | 26.49 | 34.46 | 998.27 | 34.7 | -0.24 | -1.74 | -8.21 | 33.8374 | 996.5235 | 26.7228 | 0.38263 | 0.0065238 | -0.23281
34.44 | 996.52 | 26.18 | 34.63 | 998.24 | 34.55 | -0.19 | -1.72 | -8.37 | 34.0395 | 996.4935 | 26.5728 | 0.40046 | 0.026524 | -0.39281
34.62 | 996.52 | 25.93 | 34.79 | 998.27 | 34.34 | -0.17 | -1.75 | -8.41 | 34.2298 | 996.5235 | 26.3628 | 0.39018 | -0.0034762 | -0.43281
34.77 | 996.52 | 25.78 | 34.91 | 998.26 | 34.24 | -0.14 | -1.74 | -8.46 | 34.3725 | 996.5135 | 26.2628 | 0.39747 | 0.0065238 | -0.48281
34.9 | 996.5 | 25.62 | 35.05 | 998.22 | 34.1 | -0.15 | -1.72 | -8.48 | 34.539 | 996.4735 | 26.1228 | 0.36097 | 0.026524 | -0.50281
35.04 | 996.5 | 25.49 | 35.14 | 998.22 | 34.03 | -0.1 | -1.72 | -8.54 | 34.6461 | 996.4735 | 26.0528 | 0.39394 | 0.026524 | -0.56281
35.14 | 996.49 | 25.38 | 35.26 | 998.22 | 33.97 | -0.12 | -1.73 | -8.59 | 34.7888 | 996.4735 | 25.9928 | 0.35123 | 0.016524 | -0.61281
35.26 | 996.5 | 25.31 | 35.37 | 998.23 | 33.85 | -0.11 | -1.73 | -8.54 | 34.9196 | 996.4835 | 25.8728 | 0.34041 | 0.016524 | -0.56281
35.34 | 996.48 | 25.28 | 35.42 | 998.25 | 33.84 | -0.08 | -1.77 | -8.56 | 34.9791 | 996.5035 | 25.8628 | 0.36095 | -0.023476 | -0.58281
35.41 | 996.48 | 25.16 | 35.52 | 998.25 | 33.74 | -0.11 | -1.77 | -8.58 | 35.098 | 996.5035 | 25.7628 | 0.31202 | -0.023476 | -0.60281
35.51 | 996.49 | 25.12 | 35.6 | 998.26 | 33.68 | -0.09 | -1.77 | -8.56 | 35.1931 | 996.5135 | 25.7028 | 0.31688 | -0.023476 | -0.58281
35.56 | 996.46 | 25.11 | 35.66 | 998.23 | 33.7 | -0.1 | -1.77 | -8.59 | 35.2645 | 996.4835 | 25.7228 | 0.29553 | -0.023476 | -0.61281
35.64 | 996.47 | 25.06 | 35.74 | 998.2 | 33.7 | -0.1 | -1.73 | -8.64 | 35.3596 | 996.4535 | 25.7228 | 0.28039 | 0.016524 | -0.66281
35.71 | 996.45 | 25.24 | 35.84 | 998.18 | 33.71 | -0.13 | -1.73 | -8.47 | 35.4785 | 996.4335 | 25.7328 | 0.23146 | 0.016524 | -0.49281
35.72 | 996.46 | 25.63 | 35.77 | 998.19 | 34.45 | -0.05 | -1.73 | -8.82 | 35.3953 | 996.4435 | 26.4728 | 0.32471 | 0.016524 | -0.84281
35.69 | 996.47 | 26.42 | 35.76 | 998.2 | 34.98 | -0.07 | -1.73 | -8.56 | 35.3834 | 996.4535 | 27.0028 | 0.3066 | 0.016524 | -0.58281
35.61 | 996.46 | 25.79 | 35.61 | 998.22 | 34.89 | 0 | -1.76 | -9.1 | 35.205 | 996.4735 | 26.9128 | 0.40499 | -0.013476 | -1.1228
35.15 | 996.45 | 26.45 | 35.36 | 998.21 | 35.17 | -0.21 | -1.76 | -8.72 | 34.9077 | 996.4635 | 27.1928 | 0.2423 | -0.013476 | -0.74281
34.74 | 996.43 | 27.1 | 35.2 | 998.2 | 35.45 | -0.46 | -1.77 | -8.35 | 34.7174 | 996.4535 | 27.4728 | 0.022582 | -0.023476 | -0.37281
34.49 | 996.43 | 27.33 | 34.94 | 998.16 | 35.75 | -0.45 | -1.73 | -8.42 | 34.4082 | 996.4135 | 27.7728 | 0.081788 | 0.016524 | -0.44281
33.95 | 996.42 | 27.86 | 34.62 | 998.16 | 36.12 | -0.67 | -1.74 | -8.26 | 34.0277 | 996.4135 | 28.1428 | -0.07765 | 0.0065238 | -0.28281
33.55 | 996.43 | 28.25 | 34.43 | 998.17 | 36.15 | -0.88 | -1.74 | -7.9 | 33.8017 | 996.4235 | 28.1728 | -0.25169 | 0.0065238 | 0.07719
33.23 | 996.43 | 28.5 | 34.1 | 998.18 | 36.41 | -0.87 | -1.75 | -7.91 | 33.4092 | 996.4335 | 28.4328 | -0.17924 | -0.0034762 | 0.06719
32.88 | 996.45 | 28.84 | 33.92 | 998.18 | 36.57 | -1.04 | -1.73 | -7.73 | 33.1952 | 996.4335 | 28.5928 | -0.31517 | 0.016524 | 0.24719
32.64 | 996.45 | 29.19 | 33.65 | 998.19 | 36.77 | -1.01 | -1.74 | -7.58 | 32.8741 | 996.4435 | 28.7928 | -0.23407 | 0.0065238 | 0.39719
32.43 | 996.46 | 29.37 | 33.5 | 998.17 | 36.85 | -1.07 | -1.71 | -7.48 | 32.6957 | 996.4235 | 28.8728 | -0.26569 | 0.036524 | 0.49719
32.09 | 996.47 | 29.72 | 33.21 | 998.17 | 37.2 | -1.12 | -1.7 | -7.48 | 32.3508 | 996.4235 | 29.2228 | -0.2608 | 0.046524 | 0.49719
31.76 | 996.47 | 30.06 | 33.08 | 998.18 | 37.25 | -1.32 | -1.71 | -7.19 | 32.1962 | 996.4335 | 29.2728 | -0.4362 | 0.036524 | 0.78719
31.6 | 996.46 | 30.28 | 32.85 | 998.18 | 37.32 | -1.25 | -1.72 | -7.04 | 31.9227 | 996.4335 | 29.3428 | -0.32267 | 0.026524 | 0.93719
31.5 | 996.46 | 30.49 | 32.8 | 998.18 | 37.54 | -1.3 | -1.72 | -7.05 | 31.8632 | 996.4335 | 29.5628 | -0.36321 | 0.026524 | 0.92719
31.57 | 996.46 | 30.64 | 32.89 | 998.19 | 37.33 | -1.32 | -1.73 | -6.69 | 31.9702 | 996.4435 | 29.3528 | -0.40024 | 0.016524 | 1.2872
31.56 | 996.48 | 30.69 | 32.8 | 998.2 | 37.78 | -1.24 | -1.72 | -7.09 | 31.8632 | 996.4535 | 29.8028 | -0.30321 | 0.026524 | 0.88719
31.57 | 996.49 | 30.63 | 32.78 | 998.21 | 37.37 | -1.21 | -1.72 | -6.74 | 31.8394 | 996.4635 | 29.3928 | -0.26942 | 0.026524 | 1.2372
31.43 | 996.47 | 30.69 | 32.48 | 998.21 | 37.78 | -1.05 | -1.74 | -7.09 | 31.4826 | 996.4635 | 29.8028 | -0.052647 | 0.0065238 | 0.88719
31.21 | 996.49 | 30.9 | 32.52 | 998.24 | 37.98 | -1.31 | -1.75 | -7.08 | 31.5302 | 996.4935 | 30.0028 | -0.32022 | -0.0034762 | 0.89719
31.05 | 996.48 | 30.97 | 32.22 | 998.24 | 38.09 | -1.17 | -1.76 | -7.12 | 31.1734 | 996.4935 | 30.1128 | -0.12344 | -0.013476 | 0.85719
31.02 | 996.47 | 31.05 | 32.24 | 998.2 | 37.98 | -1.22 | -1.73 | -6.93 | 31.1972 | 996.4535 | 30.0028 | -0.17723 | 0.016524 | 1.0472
30.98 | 996.47 | 31.29 | 32.22 | 998.2 | 38.28 | -1.24 | -1.73 | -6.99 | 31.1734 | 996.4535 | 30.3028 | -0.19344 | 0.016524 | 0.98719
31.06 | 996.47 | 31.18 | 32.32 | 998.2 | 37.97 | -1.26 | -1.73 | -6.79 | 31.2924 | 996.4535 | 29.9928 | -0.23237 | 0.016524 | 1.1872
31.17 | 996.46 | 31.03 | 32.25 | 998.21 | 37.93 | -1.08 | -1.75 | -6.9 | 31.2091 | 996.4635 | 29.9528 | -0.039119 | -0.0034762 | 1.0772
31.07 | 996.45 | 30.88 | 32.11 | 998.17 | 38.04 | -1.04 | -1.72 | -7.16 | 31.0426 | 996.4235 | 30.0628 | 0.027377 | 0.026524 | 0.81719
30.93 | 996.44 | 31.12 | 32.06 | 998.18 | 38.03 | -1.13 | -1.74 | -6.91 | 30.9832 | 996.4335 | 30.0528 | -0.053161 | 0.0065238 | 1.0672
30.84 | 996.43 | 31.09 | 31.99 | 998.17 | 38.05 | -1.15 | -1.74 | -6.96 | 30.8999 | 996.4235 | 30.0728 | -0.059913 | 0.0065238 | 1.0172
30.72 | 996.44 | 31.17 | 31.99 | 998.19 | 38.13 | -1.27 | -1.75 | -6.96 | 30.8999 | 996.4435 | 30.1528 | -0.17991 | -0.0034762 | 1.0172
30.65 | 996.45 | 31.22 | 31.87 | 998.2 | 38.17 | -1.22 | -1.75 | -6.95 | 30.7572 | 996.4535 | 30.1928 | -0.1072 | -0.0034762 | 1.0272
30.6 | 996.45 | 31.26 | 31.83 | 998.21 | 38.06 | -1.23 | -1.76 | -6.8 | 30.7096 | 996.4635 | 30.0828 | -0.10963 | -0.013476 | 1.1772
30.52 | 996.46 | 31.31 | 31.7 | 998.22 | 38.26 | -1.18 | -1.76 | -6.95 | 30.555 | 996.4735 | 30.2828 | -0.035029 | -0.013476 | 1.0272
30.5 | 996.46 | 31.39 | 31.68 | 998.24 | 38.22 | -1.18 | -1.78 | -6.83 | 30.5312 | 996.4935 | 30.2428 | -0.031244 | -0.033476 | 1.1472
30.49 | 996.47 | 31.35 | 31.63 | 998.24 | 38.32 | -1.14 | -1.77 | -6.97 | 30.4718 | 996.4935 | 30.3428 | 0.018219 | -0.023476 | 1.0072
30.59 | 996.47 | 31.49 | 31.95 | 998.23 | 37.86 | -1.36 | -1.76 | -6.37 | 30.8523 | 996.4835 | 29.8828 | -0.26234 | -0.013476 | 1.6072
31.23 | 996.47 | 31.2 | 32.26 | 998.22 | 37.82 | -1.03 | -1.75 | -6.62 | 31.221 | 996.4735 | 29.8428 | 0.0089887 | -0.0034762 | 1.3572
31.77 | 996.49 | 30.38 | 32.54 | 998.22 | 37.42 | -0.77 | -1.73 | -7.04 | 31.554 | 996.4735 | 29.4428 | 0.216 | 0.016524 | 0.93719
32.19 | 996.5 | 29.71 | 32.87 | 998.23 | 37.05 | -0.68 | -1.73 | -7.34 | 31.9465 | 996.4835 | 29.0728 | 0.24354 | 0.016524 | 0.63719
32.6 | 996.5 | 29.12 | 33.17 | 998.22 | 36.58 | -0.57 | -1.72 | -7.46 | 32.3032 | 996.4735 | 28.6028 | 0.29677 | 0.026524 | 0.51719
32.97 | 996.5 | 28.59 | 33.47 | 998.23 | 36.28 | -0.5 | -1.73 | -7.69 | 32.66 | 996.4835 | 28.3028 | 0.30999 | 0.016524 | 0.28719
33.3 | 996.49 | 28.18 | 33.74 | 998.2 | 35.98 | -0.44 | -1.71 | -7.8 | 32.9811 | 996.4535 | 28.0028 | 0.31889 | 0.036524 | 0.17719
33.63 | 996.48 | 27.75 | 34.02 | 998.2 | 35.65 | -0.39 | -1.72 | -7.9 | 33.3141 | 996.4535 | 27.6728 | 0.3159 | 0.026524 | 0.07719
33.89 | 996.46 | 27.36 | 34.19 | 998.18 | 35.41 | -0.3 | -1.72 | -8.05 | 33.5163 | 996.4335 | 27.4328 | 0.37373 | 0.026524 | -0.07281
34.11 | 996.48 | 27.09 | 34.4 | 998.21 | 35.08 | -0.29 | -1.73 | -7.99 | 33.766 | 996.4635 | 27.1028 | 0.34399 | 0.016524 | -0.01281
34.35 | 996.48 | 26.81 | 34.57 | 998.19 | 35.07 | -0.22 | -1.71 | -8.26 | 33.9682 | 996.4435 | 27.0928 | 0.38181 | 0.036524 | -0.28281
34.52 | 996.48 | 26.62 | 34.74 | 998.21 | 34.81 | -0.22 | -1.73 | -8.19 | 34.1704 | 996.4635 | 26.8328 | 0.34964 | 0.016524 | -0.21281
34.65 | 996.49 | 26.42 | 34.87 | 998.21 | 34.64 | -0.22 | -1.72 | -8.22 | 34.325 | 996.4635 | 26.6628 | 0.32504 | 0.026524 | -0.24281
34.79 | 996.49 | 26.26 | 34.97 | 998.2 | 34.53 | -0.18 | -1.71 | -8.27 | 34.4439 | 996.4535 | 26.5528 | 0.34611 | 0.036524 | -0.29281
34.93 | 996.48 | 26.06 | 35.09 | 998.17 | 34.48 | -0.16 | -1.69 | -8.42 | 34.5866 | 996.4235 | 26.5028 | 0.3434 | 0.056524 | -0.44281
35.04 | 996.48 | 25.95 | 35.22 | 998.17 | 34.32 | -0.18 | -1.69 | -8.37 | 34.7412 | 996.4235 | 26.3428 | 0.2988 | 0.056524 | -0.39281
35.16 | 996.48 | 25.8 | 35.31 | 998.21 | 34.25 | -0.15 | -1.73 | -8.45 | 34.8482 | 996.4635 | 26.2728 | 0.31176 | 0.016524 | -0.47281
35.24 | 996.49 | 25.74 | 35.36 | 998.23 | 34.21 | -0.12 | -1.74 | -8.47 | 34.9077 | 996.4835 | 26.2328 | 0.3323 | 0.0065238 | -0.49281
35.31 | 996.49 | 25.62 | 35.45 | 998.23 | 34.16 | -0.14 | -1.74 | -8.54 | 35.0147 | 996.4835 | 26.1828 | 0.29527 | 0.0065238 | -0.56281
35.39 | 996.48 | 25.57 | 35.51 | 998.2 | 34.03 | -0.12 | -1.72 | -8.46 | 35.0861 | 996.4535 | 26.0528 | 0.30391 | 0.026524 | -0.48281
35.46 | 996.49 | 25.55 | 35.59 | 998.23 | 33.98 | -0.13 | -1.74 | -8.43 | 35.1812 | 996.4835 | 26.0028 | 0.27877 | 0.0065238 | -0.45281
35.55 | 996.48 | 25.43 | 35.68 | 998.22 | 33.92 | -0.13 | -1.74 | -8.49 | 35.2883 | 996.4735 | 25.9428 | 0.26174 | 0.0065238 | -0.51281
35.59 | 996.48 | 25.39 | 35.74 | 998.21 | 33.86 | -0.15 | -1.73 | -8.47 | 35.3596 | 996.4635 | 25.8828 | 0.23039 | 0.016524 | -0.49281
35.67 | 996.48 | 25.33 | 35.82 | 998.19 | 33.86 | -0.15 | -1.71 | -8.53 | 35.4548 | 996.4435 | 25.8828 | 0.21524 | 0.036524 | -0.55281
35.73 | 996.46 | 25.32 | 35.85 | 998.17 | 33.82 | -0.12 | -1.71 | -8.5 | 35.4904 | 996.4235 | 25.8428 | 0.23957 | 0.036524 | -0.52281
35.77 | 996.45 | 25.26 | 35.94 | 998.17 | 33.79 | -0.17 | -1.72 | -8.53 | 35.5975 | 996.4235 | 25.8128 | 0.17253 | 0.026524 | -0.55281
35.86 | 996.44 | 25.29 | 35.98 | 998.16 | 33.79 | -0.12 | -1.72 | -8.5 | 35.645 | 996.4135 | 25.8128 | 0.21496 | 0.026524 | -0.52281
35.9 | 996.44 | 25.2 | 35.98 | 998.17 | 33.77 | -0.08 | -1.73 | -8.57 | 35.645 | 996.4235 | 25.7928 | 0.25496 | 0.016524 | -0.59281
35.95 | 996.48 | 25.17 | 36.1 | 998.21 | 33.68 | -0.15 | -1.73 | -8.51 | 35.7877 | 996.4635 | 25.7028 | 0.16225 | 0.016524 | -0.53281
36.02 | 996.46 | 25.18 | 36.16 | 998.19 | 33.74 | -0.14 | -1.73 | -8.56 | 35.8591 | 996.4435 | 25.7628 | 0.1609 | 0.016524 | -0.58281
36.07 | 996.46 | 25.08 | 36.19 | 998.18 | 33.67 | -0.12 | -1.72 | -8.59 | 35.8948 | 996.4335 | 25.6928 | 0.17522 | 0.026524 | -0.61281
36.11 | 996.45 | 25.04 | 36.21 | 998.16 | 33.6 | -0.1 | -1.71 | -8.56 | 35.9186 | 996.4135 | 25.6228 | 0.19144 | 0.036524 | -0.58281
36.16 | 996.46 | 25.03 | 36.27 | 998.17 | 33.61 | -0.11 | -1.71 | -8.58 | 35.9899 | 996.4235 | 25.6328 | 0.17008 | 0.036524 | -0.60281
36.21 | 996.47 | 25.01 | 36.34 | 998.19 | 33.6 | -0.13 | -1.72 | -8.59 | 36.0732 | 996.4435 | 25.6228 | 0.13683 | 0.026524 | -0.61281
36.26 | 996.46 | 25.01 | 36.36 | 998.17 | 33.57 | -0.1 | -1.71 | -8.56 | 36.097 | 996.4235 | 25.5928 | 0.16305 | 0.036524 | -0.58281
36.28 | 996.44 | 24.96 | 36.38 | 998.15 | 33.64 | -0.1 | -1.71 | -8.68 | 36.1207 | 996.4035 | 25.6628 | 0.15926 | 0.036524 | -0.70281
36.31 | 996.42 | 24.96 | 36.46 | 998.15 | 33.51 | -0.15 | -1.73 | -8.55 | 36.2159 | 996.4035 | 25.5328 | 0.094122 | 0.016524 | -0.57281
36.38 | 996.43 | 24.9 | 36.5 | 998.15 | 33.55 | -0.12 | -1.72 | -8.65 | 36.2634 | 996.4035 | 25.5728 | 0.11655 | 0.026524 | -0.67281
36.38 | 996.42 | 24.94 | 36.48 | 998.14 | 33.57 | -0.1 | -1.72 | -8.63 | 36.2397 | 996.3935 | 25.5928 | 0.14034 | 0.026524 | -0.65281
36.4 | 996.41 | 24.93 | 36.55 | 998.14 | 33.48 | -0.15 | -1.73 | -8.55 | 36.3229 | 996.3935 | 25.5028 | 0.07709 | 0.016524 | -0.57281
36.47 | 996.43 | 24.98 | 36.6 | 998.16 | 33.46 | -0.13 | -1.73 | -8.48 | 36.3824 | 996.4135 | 25.4828 | 0.087627 | 0.016524 | -0.50281
[5 empty] | SLOPE | 1.1893 | 0.90871 | 1.3119
[5 empty] | INTERCEPT | -7.1443 | 89.3925 | -19.3431 |  |  |  | Teplota [C] | Tlak [hPa] | Vlhkosť [%]
[5 empty] | AVG | -0.59767 | -1.7465 | -7.9772 |  |  | AVG | -3.2144e-15 | -5.6302e-14 | -4.0095e-15
[5 empty] | RMSE | 0.76209 | 1.7467 | 8.0125 |  |  | RMSE | 0.25393 | 0.022966 | 0.75157
[5 empty] | STD DEV | 0.47284 | 0.022966 | 0.75157 |  |  | STD DEV | 0.25393 | 0.022966 | 0.75157
This chart isn’t available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
This chart isn’t available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.

## validation_GML_2026_001
etalon_pressure | etalon_temperature | etalon_humidity | logger_temperature | logger_humidity | logger_pressure | logger_temperature_raw | logger_humidity_raw | logger_pressure_raw | temperature_scale | temperature_offset | humidity_scale | humidity_offset | pressure_scale | pressure_offset | validation_delta_temperature | validation_delta_pressure | validation_delta_humidity
995.87 | 36.42 | 28.21 | 37.3 | 30.76 | 995.82 | 37.374 | 38.738 | 997.566 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.88 | 0.05 | -2.55
995.86 | 36.42 | 29.03 | 37.3 | 30.73 | 995.82 | 37.372 | 38.707 | 997.566 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.88 | 0.04 | -1.7
995.84 | 36.5 | 28.4 | 37.04 | 28.47 | 995.82 | 37.153 | 36.447 | 997.568 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.54 | 0.02 | -0.07
995.83 | 36.56 | 28.26 | 36.96 | 28.64 | 995.8 | 37.089 | 36.613 | 997.545 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.4 | 0.03 | -0.38
995.81 | 36.56 | 28.51 | 36.95 | 28.8 | 995.77 | 37.083 | 36.773 | 997.516 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.39 | 0.04 | -0.29
995.78 | 36.58 | 28.78 | 36.88 | 28.64 | 995.75 | 37.024 | 36.618 | 997.499 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.3 | 0.03 | 0.14
995.77 | 36.23 | 28.24 | 36.82 | 28.74 | 995.73 | 36.971 | 36.721 | 997.474 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.59 | 0.04 | -0.5
995.77 | 35.81 | 28.59 | 36.48 | 28.75 | 995.72 | 36.688 | 36.723 | 997.468 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.67 | 0.05 | -0.16
995.77 | 35.21 | 28.88 | 36.01 | 28.85 | 995.75 | 36.287 | 36.824 | 997.495 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.8 | 0.02 | 0.03
995.78 | 34.63 | 29.39 | 35.36 | 29.39 | 995.76 | 35.743 | 37.37 | 997.503 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.73 | 0.02 | 0
995.77 | 34.02 | 30.06 | 34.87 | 29.64 | 995.77 | 35.332 | 37.614 | 997.513 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.85 | 0 | 0.42
995.74 | 33.63 | 30.46 | 34.45 | 29.81 | 995.77 | 34.974 | 37.788 | 997.513 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.82 | -0.03 | 0.65
995.74 | 33.16 | 30.86 | 34.29 | 29.83 | 995.74 | 34.846 | 37.811 | 997.485 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -1.13 | 0 | 1.03
995.72 | 32.77 | 31.19 | 33.7 | 30.15 | 995.71 | 34.343 | 38.123 | 997.461 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.93 | 0.01 | 1.04
995.71 | 32.5 | 31.57 | 33.35 | 30.46 | 995.68 | 34.05 | 38.438 | 997.431 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.85 | 0.03 | 1.11
995.74 | 32.23 | 31.76 | 33.16 | 30.51 | 995.69 | 33.895 | 38.491 | 997.441 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.93 | 0.05 | 1.25
995.73 | 31.99 | 31.85 | 32.83 | 30.79 | 995.71 | 33.612 | 38.768 | 997.453 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.84 | 0.02 | 1.06
995.73 | 31.78 | 31.99 | 32.58 | 30.77 | 995.69 | 33.404 | 38.747 | 997.435 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.8 | 0.04 | 1.22
995.72 | 31.65 | 32.11 | 32.41 | 30.71 | 995.68 | 33.265 | 38.688 | 997.424 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.76 | 0.04 | 1.4
995.7 | 31.47 | 32.24 | 32.22 | 30.62 | 995.67 | 33.104 | 38.6 | 997.412 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.75 | 0.03 | 1.62
995.67 | 31.28 | 32.26 | 31.99 | 30.74 | 995.64 | 32.912 | 38.717 | 997.39 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.71 | 0.03 | 1.52
995.64 | 31.11 | 32.38 | 31.69 | 31.06 | 995.64 | 32.656 | 39.041 | 997.391 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.58 | 0 | 1.32
995.64 | 31.13 | 32.31 | 31.55 | 31.06 | 995.62 | 32.538 | 39.039 | 997.362 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.42 | 0.02 | 1.25
995.62 | 31 | 32.33 | 31.64 | 30.67 | 995.6 | 32.613 | 38.648 | 997.347 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.64 | 0.02 | 1.66
995.62 | 30.95 | 32.42 | 31.32 | 30.99 | 995.57 | 32.346 | 38.972 | 997.319 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.37 | 0.05 | 1.43
995.61 | 30.94 | 32.36 | 31.33 | 30.89 | 995.57 | 32.351 | 38.869 | 997.312 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.39 | 0.04 | 1.47
995.59 | 30.9 | 32.43 | 31.14 | 31.03 | 995.56 | 32.197 | 39.005 | 997.306 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.24 | 0.03 | 1.4
995.56 | 30.97 | 32.07 | 31.23 | 30.84 | 995.55 | 32.271 | 38.814 | 997.292 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.26 | 0.01 | 1.23
995.56 | 30.94 | 32.05 | 31.14 | 30.83 | 995.53 | 32.191 | 38.812 | 997.273 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.2 | 0.03 | 1.22
995.52 | 30.77 | 32.27 | 31.07 | 30.79 | 995.51 | 32.132 | 38.77 | 997.252 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.3 | 0.01 | 1.48
995.47 | 30.74 | 32.23 | 31.01 | 30.85 | 995.47 | 32.084 | 38.825 | 997.217 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.27 | 0 | 1.38
995.47 | 30.64 | 32.31 | 30.84 | 30.9 | 995.45 | 31.946 | 38.877 | 997.196 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.2 | 0.02 | 1.41
995.49 | 30.63 | 32.3 | 30.8 | 30.96 | 995.45 | 31.908 | 38.936 | 997.194 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.17 | 0.04 | 1.34
995.44 | 30.55 | 32.41 | 30.77 | 30.99 | 995.46 | 31.881 | 38.97 | 997.207 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.22 | -0.02 | 1.42
995.41 | 30.62 | 32.24 | 30.68 | 31.05 | 995.41 | 31.807 | 39.026 | 997.161 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.06 | 0 | 1.19
995.41 | 30.56 | 32.19 | 30.7 | 30.92 | 995.39 | 31.823 | 38.902 | 997.137 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.14 | 0.02 | 1.27
995.4 | 30.46 | 32.31 | 30.51 | 31.13 | 995.37 | 31.668 | 39.111 | 997.121 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.05 | 0.03 | 1.18
995.39 | 30.4 | 32.34 | 30.48 | 31 | 995.38 | 31.641 | 38.976 | 997.129 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.08 | 0.01 | 1.34
995.44 | 30.34 | 32.34 | 30.46 | 30.98 | 995.38 | 31.625 | 38.953 | 997.129 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.12 | 0.06 | 1.36
995.4 | 30.27 | 32.34 | 30.34 | 30.99 | 995.43 | 31.518 | 38.97 | 997.176 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.07 | -0.03 | 1.35
995.34 | 30.36 | 33.01 | 30.34 | 31.03 | 995.37 | 31.518 | 39.01 | 997.116 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.02 | -0.03 | 1.98
995.34 | 30.47 | 32.46 | 30.49 | 31.25 | 995.33 | 31.652 | 39.23 | 997.076 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.02 | 0.01 | 1.21
995.34 | 30.34 | 32.55 | 30.34 | 31.39 | 995.33 | 31.518 | 39.363 | 997.08 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0 | 0.01 | 1.16
995.3 | 30.26 | 32.6 | 30.22 | 31.38 | 995.33 | 31.422 | 39.354 | 997.075 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | -0.03 | 1.22
995.29 | 30.32 | 32.41 | 30.18 | 31.4 | 995.3 | 31.39 | 39.38 | 997.048 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | -0.01 | 1.01
995.28 | 30.17 | 32.47 | 30.25 | 31.1 | 995.27 | 31.444 | 39.075 | 997.017 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.08 | 0.01 | 1.37
995.25 | 30.14 | 32.5 | 30.11 | 31.28 | 995.28 | 31.326 | 39.256 | 997.023 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.03 | -0.03 | 1.22
995.26 | 30.07 | 32.47 | 30.02 | 31.16 | 995.27 | 31.257 | 39.132 | 997.013 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | -0.01 | 1.31
995.22 | 30 | 32.59 | 29.91 | 31.38 | 995.25 | 31.16 | 39.357 | 997.001 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | -0.03 | 1.21
995.18 | 29.99 | 32.59 | 29.97 | 31.24 | 995.23 | 31.209 | 39.216 | 996.975 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.02 | -0.05 | 1.35
995.21 | 29.99 | 32.48 | 30.03 | 31.19 | 995.2 | 31.262 | 39.163 | 996.942 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.04 | 0.01 | 1.29
995.2 | 29.97 | 32.43 | 29.94 | 31.09 | 995.22 | 31.187 | 39.066 | 996.962 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.03 | -0.02 | 1.34
995.19 | 29.97 | 32.38 | 29.96 | 31.06 | 995.21 | 31.203 | 39.037 | 996.958 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.01 | -0.02 | 1.32
995.2 | 29.87 | 32.52 | 29.71 | 31.23 | 995.21 | 30.995 | 39.211 | 996.958 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | -0.01 | 1.29
995.21 | 29.82 | 32.49 | 29.89 | 31.08 | 995.22 | 31.139 | 39.06 | 996.966 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.07 | -0.01 | 1.41
995.17 | 29.88 | 33.02 | 29.64 | 31.45 | 995.21 | 30.936 | 39.432 | 996.957 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | -0.04 | 1.57
995.17 | 30.29 | 32.41 | 29.98 | 31.25 | 995.21 | 31.219 | 39.226 | 996.958 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.31 | -0.04 | 1.16
995.19 | 30.8 | 31.75 | 30.22 | 30.93 | 995.18 | 31.417 | 38.907 | 996.925 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.58 | 0.01 | 0.82
995.13 | 31.27 | 31.03 | 30.55 | 30.54 | 995.18 | 31.7 | 38.518 | 996.925 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.72 | -0.05 | 0.49
995.1 | 31.67 | 30.39 | 30.93 | 30.03 | 995.1 | 32.02 | 38.011 | 996.847 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.74 | 0 | 0.36
995.09 | 32.05 | 29.83 | 31.29 | 29.61 | 995.05 | 32.319 | 37.585 | 996.798 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.76 | 0.04 | 0.22
995.08 | 32.36 | 29.33 | 31.61 | 29.22 | 995.06 | 32.586 | 37.198 | 996.803 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.75 | 0.02 | 0.11
995.06 | 32.65 | 28.92 | 31.89 | 28.87 | 995.04 | 32.821 | 36.851 | 996.791 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.76 | 0.02 | 0.05
995.05 | 32.91 | 28.55 | 32.15 | 28.54 | 995.03 | 33.046 | 36.514 | 996.773 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.76 | 0.02 | 0.01
995.09 | 33.13 | 28.17 | 32.41 | 28.31 | 995.02 | 33.265 | 36.285 | 996.763 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.72 | 0.07 | -0.14
995.08 | 33.34 | 27.88 | 32.64 | 28.01 | 995.04 | 33.457 | 35.987 | 996.787 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.7 | 0.04 | -0.13
995.05 | 33.54 | 27.64 | 32.81 | 27.83 | 995.06 | 33.601 | 35.81 | 996.81 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.73 | -0.01 | -0.19
995.03 | 33.7 | 27.41 | 32.99 | 27.66 | 995.03 | 33.751 | 35.638 | 996.78 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.71 | 0 | -0.25
994.99 | 33.83 | 27.19 | 33.18 | 27.53 | 995 | 33.906 | 35.506 | 996.742 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.65 | -0.01 | -0.34
994.97 | 33.95 | 27.07 | 33.31 | 27.35 | 994.97 | 34.023 | 35.327 | 996.717 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.64 | 0 | -0.28
994.96 | 34.05 | 26.92 | 33.41 | 27.29 | 994.94 | 34.103 | 35.264 | 996.688 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.64 | 0.02 | -0.37
994.92 | 34.18 | 26.79 | 33.57 | 27.17 | 994.92 | 34.237 | 35.146 | 996.67 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.61 | 0 | -0.38
994.88 | 34.26 | 26.67 | 33.66 | 27.06 | 994.91 | 34.317 | 35.037 | 996.661 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.6 | -0.03 | -0.39
994.88 | 34.35 | 26.59 | 33.73 | 27.01 | 994.88 | 34.376 | 34.991 | 996.622 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.62 | 0 | -0.42
994.87 | 34.43 | 26.48 | 33.86 | 26.87 | 994.84 | 34.482 | 34.845 | 996.591 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.57 | 0.03 | -0.39
994.84 | 34.5 | 26.42 | 33.97 | 26.82 | 994.85 | 34.573 | 34.797 | 996.601 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.53 | -0.01 | -0.4
994.8 | 34.55 | 26.36 | 34.01 | 26.86 | 994.81 | 34.611 | 34.835 | 996.557 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.54 | -0.01 | -0.5
994.76 | 34.61 | 26.31 | 34.09 | 26.75 | 994.79 | 34.675 | 34.723 | 996.534 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.52 | -0.03 | -0.44
994.76 | 34.68 | 26.23 | 34.15 | 26.71 | 994.74 | 34.728 | 34.692 | 996.485 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.53 | 0.02 | -0.48
994.81 | 34.73 | 26.19 | 34.2 | 26.63 | 994.73 | 34.771 | 34.608 | 996.481 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.53 | 0.08 | -0.44
994.78 | 34.78 | 26.15 | 34.3 | 26.57 | 994.79 | 34.851 | 34.551 | 996.533 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.48 | -0.01 | -0.42
994.75 | 34.82 | 26.11 | 34.33 | 26.59 | 994.75 | 34.878 | 34.562 | 996.494 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.49 | 0 | -0.48
994.75 | 34.88 | 26.09 | 34.42 | 26.5 | 994.71 | 34.952 | 34.475 | 996.453 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.46 | 0.04 | -0.41
994.72 | 34.91 | 26.25 | 34.43 | 26.52 | 994.73 | 34.958 | 34.494 | 996.475 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.48 | -0.01 | -0.27
994.69 | 34.94 | 26.72 | 34.57 | 28.39 | 994.71 | 35.075 | 36.365 | 996.457 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.37 | -0.02 | -1.67
994.69 | 34.91 | 27.04 | 34.49 | 27.17 | 994.69 | 35.011 | 35.15 | 996.436 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.42 | 0 | -0.13
994.68 | 34.91 | 27.16 | 34.5 | 27.59 | 994.68 | 35.016 | 35.569 | 996.424 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.41 | 0 | -0.43
994.66 | 34.96 | 26.99 | 34.59 | 27.49 | 994.67 | 35.091 | 35.47 | 996.414 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.37 | -0.01 | -0.5
994.66 | 35.03 | 26.82 | 34.67 | 27.3 | 994.65 | 35.166 | 35.281 | 996.401 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.36 | 0.01 | -0.48
994.63 | 35.08 | 26.72 | 34.77 | 27.14 | 994.64 | 35.246 | 35.115 | 996.387 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.31 | -0.01 | -0.42
994.65 | 35.11 | 26.65 | 34.85 | 27 | 994.65 | 35.31 | 34.98 | 996.392 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0 | -0.35
994.66 | 35.17 | 26.54 | 34.91 | 26.91 | 994.65 | 35.364 | 34.89 | 996.399 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0.01 | -0.37
994.66 | 35.22 | 26.47 | 35 | 26.77 | 994.65 | 35.444 | 34.743 | 996.399 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.01 | -0.3
994.67 | 35.29 | 26.36 | 35.09 | 26.65 | 994.65 | 35.513 | 34.623 | 996.394 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.02 | -0.29
994.67 | 35.32 | 26.35 | 35.15 | 26.62 | 994.65 | 35.566 | 34.6 | 996.399 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.02 | -0.27
994.63 | 35.34 | 26.34 | 35.16 | 26.61 | 994.62 | 35.572 | 34.591 | 996.368 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.01 | -0.27
994.57 | 35.37 | 26.32 | 35.25 | 26.58 | 994.58 | 35.647 | 34.553 | 996.331 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | -0.01 | -0.26
994.67 | 35.38 | 26.34 | 35.28 | 26.55 | 994.55 | 35.679 | 34.526 | 996.294 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.12 | -0.21
994.62 | 35.4 | 26.32 | 35.28 | 26.59 | 994.64 | 35.679 | 34.566 | 996.39 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | -0.02 | -0.27
994.65 | 35.46 | 26.24 | 35.34 | 26.53 | 994.58 | 35.727 | 34.511 | 996.329 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.07 | -0.29
994.6 | 35.49 | 26.21 | 35.44 | 26.45 | 994.61 | 35.812 | 34.425 | 996.354 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | -0.01 | -0.24
994.6 | 35.54 | 26.19 | 35.44 | 26.4 | 994.55 | 35.807 | 34.377 | 996.297 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.05 | -0.21
994.61 | 35.54 | 26.24 | 35.5 | 26.4 | 994.58 | 35.86 | 34.375 | 996.322 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | 0.03 | -0.16
994.59 | 35.56 | 26.2 | 35.51 | 26.41 | 994.58 | 35.866 | 34.389 | 996.326 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0.01 | -0.21
994.64 | 35.57 | 26.24 | 35.49 | 26.42 | 994.56 | 35.855 | 34.396 | 996.306 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.08 | -0.18
994.61 | 35.59 | 26.21 | 35.56 | 26.47 | 994.61 | 35.908 | 34.444 | 996.358 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.03 | 0 | -0.26
994.58 | 35.63 | 26.13 | 35.6 | 26.3 | 994.56 | 35.946 | 34.276 | 996.306 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.03 | 0.02 | -0.17
994.56 | 35.64 | 26.16 | 35.69 | 26.31 | 994.52 | 36.02 | 34.291 | 996.264 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.05 | 0.04 | -0.15
994.55 | 35.66 | 26.1 | 35.65 | 26.3 | 994.49 | 35.983 | 34.28 | 996.236 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.01 | 0.06 | -0.2
994.56 | 35.69 | 26.13 | 35.73 | 26.32 | 994.5 | 36.058 | 34.297 | 996.249 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.04 | 0.06 | -0.19
994.57 | 35.71 | 26.11 | 35.73 | 26.34 | 994.52 | 36.052 | 34.314 | 996.264 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.02 | 0.05 | -0.23
994.52 | 35.72 | 26.15 | 35.75 | 26.3 | 994.52 | 36.074 | 34.276 | 996.262 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.03 | 0 | -0.15
994.51 | 35.74 | 26.11 | 35.77 | 26.3 | 994.46 | 36.09 | 34.282 | 996.209 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.03 | 0.05 | -0.19
994.45 | 35.74 | 26.13 | 35.82 | 26.29 | 994.45 | 36.133 | 34.263 | 996.194 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.08 | 0 | -0.16
994.46 | 35.76 | 26.1 | 35.79 | 26.27 | 994.38 | 36.101 | 34.249 | 996.131 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.03 | 0.08 | -0.17
994.47 | 35.78 | 26.08 | 35.8 | 26.24 | 994.39 | 36.117 | 34.219 | 996.138 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.02 | 0.08 | -0.16
994.48 | 35.8 | 26.08 | 35.87 | 26.23 | 994.41 | 36.17 | 34.211 | 996.155 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.07 | 0.07 | -0.15
994.47 | 35.81 | 26.06 | 35.89 | 26.24 | 994.44 | 36.191 | 34.213 | 996.186 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.08 | 0.03 | -0.18
994.48 | 35.84 | 26.08 | 35.9 | 26.2 | 994.43 | 36.197 | 34.177 | 996.173 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.06 | 0.05 | -0.12
994.49 | 35.84 | 26.07 | 35.93 | 26.21 | 994.42 | 36.218 | 34.19 | 996.171 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.09 | 0.07 | -0.14
994.46 | 35.87 | 26.08 | 35.94 | 26.27 | 994.43 | 36.229 | 34.246 | 996.176 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.07 | 0.03 | -0.19
994.46 | 35.86 | 26.04 | 35.96 | 26.22 | 994.42 | 36.245 | 34.194 | 996.168 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.1 | 0.04 | -0.18
994.44 | 35.89 | 26.04 | 35.96 | 26.19 | 994.4 | 36.245 | 34.171 | 996.144 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.07 | 0.04 | -0.15
994.47 | 35.91 | 26.02 | 36 | 26.16 | 994.39 | 36.282 | 34.139 | 996.132 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.09 | 0.08 | -0.14
994.44 | 35.93 | 26.02 | 36.03 | 26.09 | 994.44 | 36.309 | 34.07 | 996.183 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.1 | 0 | -0.07
994.41 | 35.93 | 26.07 | 36.01 | 26.19 | 994.39 | 36.287 | 34.169 | 996.141 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.08 | 0.02 | -0.12
994.4 | 35.93 | 26.07 | 36.03 | 26.19 | 994.36 | 36.309 | 34.171 | 996.108 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.1 | 0.04 | -0.12
994.37 | 35.97 | 26.02 | 36.07 | 26.18 | 994.35 | 36.341 | 34.162 | 996.098 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.1 | 0.02 | -0.16
994.35 | 35.97 | 26.02 | 36.09 | 26.12 | 994.33 | 36.357 | 34.101 | 996.075 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.12 | 0.02 | -0.1
994.36 | 36.01 | 26.02 | 36.12 | 26.11 | 994.3 | 36.378 | 34.091 | 996.043 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.11 | 0.06 | -0.09
994.36 | 35.98 | 26.05 | 36.13 | 26.13 | 994.31 | 36.394 | 34.105 | 996.053 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.15 | 0.05 | -0.08
994.36 | 35.98 | 26.07 | 36.12 | 26.18 | 994.3 | 36.378 | 34.158 | 996.051 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.14 | 0.06 | -0.11
994.32 | 36 | 26.06 | 36.12 | 26.13 | 994.31 | 36.384 | 34.106 | 996.058 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.12 | 0.01 | -0.07
994.39 | 36.04 | 26.03 | 36.15 | 26.14 | 994.26 | 36.405 | 34.12 | 996.01 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.11 | 0.13 | -0.11
994.37 | 36.05 | 26.04 | 36.16 | 26.12 | 994.34 | 36.416 | 34.095 | 996.084 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.11 | 0.03 | -0.08
994.38 | 36.06 | 26.06 | 36.19 | 26.13 | 994.32 | 36.437 | 34.11 | 996.069 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.13 | 0.06 | -0.07
994.37 | 36.04 | 26.05 | 36.16 | 26.2 | 994.33 | 36.416 | 34.175 | 996.079 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.12 | 0.04 | -0.15
994.35 | 36.06 | 26.03 | 36.22 | 26.17 | 994.32 | 36.469 | 34.146 | 996.066 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.16 | 0.03 | -0.14
994.34 | 36.08 | 26.05 | 36.21 | 26.17 | 994.3 | 36.458 | 34.148 | 996.043 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.13 | 0.04 | -0.12
994.31 | 36.09 | 26 | 36.22 | 26.15 | 994.29 | 36.469 | 34.126 | 996.04 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.13 | 0.02 | -0.15
994.3 | 36.09 | 26.05 | 36.28 | 26.16 | 994.28 | 36.517 | 34.135 | 996.03 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.19 | 0.02 | -0.11
994.3 | 36.1 | 26.07 | 36.25 | 26.12 | 994.27 | 36.49 | 34.097 | 996.013 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.15 | 0.03 | -0.05
994.3 | 36.09 | 26.05 | 36.27 | 26.16 | 994.25 | 36.506 | 34.139 | 996 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.18 | 0.05 | -0.11
994.34 | 36.12 | 26.05 | 36.26 | 26.17 | 994.26 | 36.501 | 34.15 | 996.007 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.14 | 0.08 | -0.12
994.3 | 36.11 | 26.03 | 36.27 | 26.21 | 994.29 | 36.506 | 34.187 | 996.038 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.16 | 0.01 | -0.18
994.3 | 36.13 | 26.05 | 36.3 | 26.16 | 994.27 | 36.533 | 34.135 | 996.021 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.17 | 0.03 | -0.11
994.3 | 36.14 | 26.04 | 36.34 | 26.11 | 994.27 | 36.565 | 34.089 | 996.015 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.2 | 0.03 | -0.07
994.34 | 36.16 | 26.05 | 36.34 | 26.14 | 994.24 | 36.565 | 34.118 | 995.985 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.18 | 0.1 | -0.09
994.32 | 36.15 | 26.05 | 36.34 | 26.14 | 994.28 | 36.571 | 34.12 | 996.03 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.19 | 0.04 | -0.09
994.28 | 36.17 | 26.03 | 36.35 | 26.14 | 994.26 | 36.576 | 34.122 | 996.004 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.18 | 0.02 | -0.11
994.27 | 36.18 | 26 | 36.37 | 26.09 | 994.26 | 36.592 | 34.07 | 996.004 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.19 | 0.01 | -0.09
994.25 | 36.2 | 26.04 | 36.37 | 26.1 | 994.22 | 36.592 | 34.08 | 995.966 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.17 | 0.03 | -0.06
994.25 | 36.19 | 26.03 | 36.37 | 26.09 | 994.21 | 36.592 | 34.07 | 995.958 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.18 | 0.04 | -0.06
994.21 | 36.19 | 26.06 | 36.36 | 26.09 | 994.2 | 36.581 | 34.063 | 995.951 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.17 | 0.01 | -0.03
994.19 | 36.17 | 26.06 | 36.4 | 26.11 | 994.16 | 36.619 | 34.084 | 995.906 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.23 | 0.03 | -0.05
994.26 | 36.2 | 26.08 | 36.39 | 26.16 | 994.14 | 36.608 | 34.141 | 995.883 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.19 | 0.12 | -0.08
994.23 | 36.21 | 26.06 | 36.4 | 26.15 | 994.21 | 36.613 | 34.124 | 995.958 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.19 | 0.02 | -0.09
994.2 | 36.22 | 26.08 | 36.43 | 26.12 | 994.19 | 36.645 | 34.099 | 995.937 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.21 | 0.01 | -0.04
994.24 | 36.22 | 26.07 | 36.45 | 26.17 | 994.17 | 36.661 | 34.148 | 995.913 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.23 | 0.07 | -0.1
994.19 | 36.24 | 26.02 | 36.46 | 26.12 | 994.19 | 36.667 | 34.097 | 995.94 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.22 | 0 | -0.1
994.23 | 36.25 | 26.08 | 36.48 | 26.11 | 994.16 | 36.688 | 34.084 | 995.902 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.23 | 0.07 | -0.03
994.25 | 36.26 | 26.08 | 36.48 | 26.17 | 994.17 | 36.688 | 34.143 | 995.921 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.22 | 0.08 | -0.09
994.24 | 36.28 | 26.07 | 36.48 | 26.12 | 994.19 | 36.683 | 34.095 | 995.933 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.2 | 0.05 | -0.05
994.22 | 36.26 | 26.09 | 36.47 | 26.13 | 994.18 | 36.672 | 34.105 | 995.926 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.21 | 0.04 | -0.04
994.26 | 36.28 | 26.1 | 36.48 | 26.11 | 994.16 | 36.688 | 34.087 | 995.911 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.2 | 0.1 | -0.01
994.27 | 36.27 | 26.11 | 36.48 | 26.17 | 994.21 | 36.688 | 34.15 | 995.952 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.21 | 0.06 | -0.06
994.25 | 36.27 | 26.12 | 36.49 | 26.19 | 994.21 | 36.693 | 34.166 | 995.959 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.22 | 0.04 | -0.07
994.21 | 36.26 | 26.13 | 36.49 | 26.17 | 994.18 | 36.693 | 34.15 | 995.922 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.23 | 0.03 | -0.04
994.21 | 36.29 | 26.14 | 36.51 | 26.2 | 994.16 | 36.709 | 34.175 | 995.903 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.22 | 0.05 | -0.06
994.22 | 36.28 | 26.13 | 36.5 | 26.21 | 994.16 | 36.699 | 34.188 | 995.906 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.22 | 0.06 | -0.08
994.18 | 36.29 | 26.1 | 36.5 | 26.17 | 994.16 | 36.699 | 34.145 | 995.907 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.21 | 0.02 | -0.07
994.18 | 36.29 | 26.14 | 36.51 | 26.2 | 994.14 | 36.709 | 34.175 | 995.883 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.22 | 0.04 | -0.06
994.18 | 36.3 | 26.12 | 36.53 | 26.22 | 994.14 | 36.725 | 34.192 | 995.888 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.23 | 0.04 | -0.1
994.18 | 36.32 | 26.12 | 36.55 | 26.19 | 994.11 | 36.741 | 34.169 | 995.858 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.23 | 0.07 | -0.07
994.15 | 36.34 | 26.14 | 36.57 | 26.19 | 994.15 | 36.763 | 34.166 | 995.894 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.23 | 0 | -0.05
994.14 | 36.32 | 26.13 | 36.54 | 26.17 | 994.1 | 36.736 | 34.146 | 995.845 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.22 | 0.04 | -0.04
994.14 | 36.32 | 26.14 | 36.54 | 26.2 | 994.09 | 36.736 | 34.177 | 995.832 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.22 | 0.05 | -0.06
994.1 | 36.33 | 26.15 | 36.56 | 26.18 | 994.11 | 36.752 | 34.154 | 995.86 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.23 | -0.01 | -0.03
994.07 | 36.33 | 26.13 | 36.56 | 26.15 | 994.06 | 36.752 | 34.124 | 995.81 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.23 | 0.01 | -0.02
994.04 | 36.33 | 26.08 | 36.54 | 26.24 | 994.04 | 36.736 | 34.217 | 995.785 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.21 | 0 | -0.16
994.04 | 36.34 | 26.1 | 36.57 | 26.13 | 993.98 | 36.763 | 34.108 | 995.729 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.23 | 0.06 | -0.03
994.05 | 36.34 | 26.19 | 36.56 | 26.16 | 993.99 | 36.752 | 34.137 | 995.741 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.22 | 0.06 | 0.03
994.01 | 35.93 | 26.37 | 36.56 | 26.32 | 994.01 | 36.752 | 34.297 | 995.761 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.63 | 0 | 0.05
993.98 | 35.22 | 26.73 | 36.1 | 26.54 | 993.97 | 36.362 | 34.518 | 995.713 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.88 | 0.01 | 0.19
993.98 | 34.82 | 26.85 | 35.46 | 27.04 | 993.93 | 35.823 | 35.018 | 995.679 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.64 | 0.05 | -0.19
994.03 | 34.27 | 27.56 | 35.08 | 27.08 | 993.92 | 35.508 | 35.062 | 995.667 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.81 | 0.11 | 0.48
994.02 | 33.5 | 28.43 | 34.67 | 27.29 | 993.97 | 35.166 | 35.262 | 995.72 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -1.17 | 0.05 | 1.14
994.02 | 33.14 | 28.88 | 34.05 | 27.79 | 993.99 | 34.637 | 35.772 | 995.741 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.91 | 0.03 | 1.09
993.99 | 32.73 | 29.18 | 33.87 | 27.75 | 993.98 | 34.493 | 35.726 | 995.724 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -1.14 | 0.01 | 1.43
993.98 | 32.34 | 29.63 | 33.38 | 28.44 | 993.96 | 34.076 | 36.412 | 995.706 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -1.04 | 0.02 | 1.19
993.95 | 32.18 | 30.21 | 33.17 | 28.49 | 993.95 | 33.903 | 36.468 | 995.699 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.99 | 0 | 1.72
993.98 | 32.12 | 30.23 | 33.06 | 28.59 | 993.92 | 33.812 | 36.571 | 995.663 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.94 | 0.06 | 1.64
993.95 | 31.84 | 30.36 | 32.99 | 28.66 | 993.95 | 33.748 | 36.634 | 995.7 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -1.15 | 0 | 1.7
993.95 | 31.52 | 30.6 | 32.42 | 28.85 | 993.94 | 33.273 | 36.823 | 995.686 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.9 | 0.01 | 1.75
993.91 | 31.33 | 30.73 | 32.29 | 28.93 | 993.9 | 33.161 | 36.912 | 995.649 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.96 | 0.01 | 1.8
993.92 | 31.18 | 30.88 | 32.08 | 28.98 | 993.88 | 32.984 | 36.954 | 995.626 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.9 | 0.04 | 1.9
993.92 | 30.98 | 30.99 | 31.99 | 28.97 | 993.89 | 32.91 | 36.947 | 995.637 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -1.01 | 0.03 | 2.02
993.93 | 30.8 | 30.84 | 31.74 | 28.94 | 993.89 | 32.701 | 36.922 | 995.636 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.94 | 0.04 | 1.9
993.9 | 30.6 | 31.07 | 31.32 | 29.06 | 993.88 | 32.349 | 37.036 | 995.631 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.72 | 0.02 | 2.01
993.95 | 30.5 | 31.15 | 31.43 | 28.85 | 993.88 | 32.44 | 36.828 | 995.625 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.93 | 0.07 | 2.3
993.98 | 30.41 | 31.15 | 31.15 | 28.97 | 993.9 | 32.199 | 36.947 | 995.647 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.74 | 0.08 | 2.18
994 | 30.48 | 30.91 | 31.19 | 28.92 | 993.89 | 32.237 | 36.897 | 995.638 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.71 | 0.11 | 1.99
994.05 | 30.42 | 30.8 | 31.04 | 28.86 | 993.95 | 32.108 | 36.838 | 995.694 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.62 | 0.1 | 1.94
994.08 | 30.23 | 30.93 | 30.82 | 28.81 | 993.97 | 31.921 | 36.782 | 995.721 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.59 | 0.11 | 2.12
994.1 | 30.15 | 31.03 | 30.86 | 28.92 | 994.02 | 31.959 | 36.895 | 995.764 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.71 | 0.08 | 2.11
994.1 | 30.13 | 31.15 | 30.54 | 29.21 | 994.04 | 31.692 | 37.183 | 995.783 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.41 | 0.06 | 1.94
994.1 | 30.2 | 30.96 | 30.63 | 29.02 | 994.06 | 31.761 | 36.996 | 995.806 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.43 | 0.04 | 1.94
994.08 | 30.08 | 30.79 | 30.61 | 28.8 | 994.04 | 31.751 | 36.779 | 995.789 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.53 | 0.04 | 1.99
994.09 | 29.99 | 31.03 | 30.29 | 29.08 | 994.03 | 31.484 | 37.059 | 995.776 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.3 | 0.06 | 1.95
994.09 | 30.04 | 31.18 | 30.4 | 28.94 | 994.03 | 31.569 | 36.918 | 995.779 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.36 | 0.06 | 2.24
994.05 | 30.16 | 31.03 | 30.48 | 28.99 | 994.03 | 31.638 | 36.971 | 995.775 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.32 | 0.02 | 2.04
994.06 | 30.14 | 31.01 | 30.38 | 29.17 | 994.03 | 31.558 | 37.151 | 995.772 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.24 | 0.03 | 1.84
994.1 | 30.09 | 31.03 | 30.37 | 29.22 | 994.01 | 31.548 | 37.2 | 995.758 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.28 | 0.09 | 1.81
994.07 | 30.02 | 31.05 | 30.3 | 29.2 | 994.06 | 31.489 | 37.175 | 995.809 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.28 | 0.01 | 1.85
994.08 | 29.92 | 31.18 | 30.16 | 29.26 | 994.02 | 31.371 | 37.242 | 995.768 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.24 | 0.06 | 1.92
994.11 | 29.95 | 31.11 | 30.19 | 29.26 | 994.03 | 31.393 | 37.235 | 995.777 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.24 | 0.08 | 1.85
994.1 | 29.83 | 31.17 | 30.12 | 29.21 | 994.06 | 31.339 | 37.191 | 995.81 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.29 | 0.04 | 1.96
994.09 | 29.89 | 31.08 | 30.02 | 29.38 | 994.05 | 31.249 | 37.353 | 995.798 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.13 | 0.04 | 1.7
994.07 | 29.83 | 31.07 | 29.95 | 29.27 | 994.05 | 31.195 | 37.248 | 995.801 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.12 | 0.02 | 1.8
994.04 | 29.71 | 31.14 | 29.92 | 29.32 | 994.04 | 31.168 | 37.301 | 995.789 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.21 | 0 | 1.82
994.01 | 29.72 | 31.12 | 29.84 | 29.38 | 994.02 | 31.099 | 37.357 | 995.768 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.12 | -0.01 | 1.74
994.05 | 29.62 | 31.17 | 29.74 | 29.39 | 993.99 | 31.019 | 37.372 | 995.736 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.12 | 0.06 | 1.78
994.05 | 29.59 | 31.2 | 29.64 | 29.47 | 994.03 | 30.933 | 37.446 | 995.777 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.05 | 0.02 | 1.73
994.03 | 29.55 | 31.16 | 29.66 | 29.34 | 994.03 | 30.949 | 37.313 | 995.775 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.11 | 0 | 1.82
994.01 | 29.48 | 31.24 | 29.42 | 29.63 | 993.99 | 30.747 | 37.605 | 995.732 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0.02 | 1.61
993.99 | 29.44 | 31.22 | 29.46 | 29.5 | 993.96 | 30.779 | 37.477 | 995.71 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.02 | 0.03 | 1.72
993.99 | 29.36 | 31.3 | 29.33 | 29.6 | 993.94 | 30.672 | 37.58 | 995.69 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.03 | 0.05 | 1.7
993.96 | 29.37 | 31.28 | 29.26 | 29.64 | 993.94 | 30.613 | 37.622 | 995.685 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.02 | 1.64
993.97 | 29.35 | 31.19 | 29.41 | 29.42 | 993.91 | 30.736 | 37.397 | 995.654 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.06 | 0.06 | 1.77
993.99 | 29.32 | 31.25 | 29.28 | 29.46 | 993.92 | 30.634 | 37.437 | 995.667 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | 0.07 | 1.79
993.98 | 29.3 | 31.24 | 29.26 | 29.46 | 993.93 | 30.613 | 37.439 | 995.674 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | 0.05 | 1.78
993.96 | 29.25 | 31.28 | 29.14 | 29.53 | 993.92 | 30.512 | 37.505 | 995.665 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.04 | 1.75
993.97 | 29.29 | 31.18 | 29.18 | 29.53 | 993.9 | 30.549 | 37.503 | 995.642 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.07 | 1.65
993.95 | 29.22 | 31.24 | 29.09 | 29.52 | 993.91 | 30.469 | 37.496 | 995.655 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.04 | 1.72
993.98 | 29.2 | 31.11 | 29.08 | 29.5 | 993.92 | 30.458 | 37.481 | 995.669 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 1.61
993.95 | 29.17 | 30.93 | 28.99 | 29.46 | 993.95 | 30.389 | 37.441 | 995.701 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0 | 1.47
993.94 | 29.05 | 30.98 | 28.88 | 29.37 | 993.92 | 30.293 | 37.343 | 995.667 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.02 | 1.61
993.93 | 29.06 | 30.98 | 28.87 | 29.38 | 993.92 | 30.282 | 37.357 | 995.663 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.01 | 1.6
993.92 | 29.01 | 31.02 | 28.9 | 29.34 | 993.91 | 30.314 | 37.313 | 995.66 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.01 | 1.68
993.91 | 29.04 | 31 | 28.8 | 29.4 | 993.9 | 30.223 | 37.381 | 995.649 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.01 | 1.6
993.89 | 29.07 | 30.76 | 28.86 | 29.28 | 993.87 | 30.277 | 37.256 | 995.617 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.02 | 1.48
993.9 | 28.92 | 30.85 | 28.71 | 29.3 | 993.88 | 30.148 | 37.273 | 995.627 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.02 | 1.55
993.92 | 28.9 | 30.97 | 28.65 | 29.31 | 993.9 | 30.1 | 37.292 | 995.647 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.02 | 1.66
993.9 | 28.85 | 31.01 | 28.49 | 29.57 | 993.91 | 29.967 | 37.547 | 995.659 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.36 | -0.01 | 1.44
993.88 | 28.91 | 30.8 | 28.59 | 29.35 | 993.89 | 30.047 | 37.332 | 995.638 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.32 | -0.01 | 1.45
993.84 | 28.83 | 30.79 | 28.57 | 29.08 | 993.88 | 30.036 | 37.057 | 995.625 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | -0.04 | 1.71
993.81 | 28.84 | 30.81 | 28.47 | 29.21 | 993.82 | 29.945 | 37.189 | 995.566 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.37 | -0.01 | 1.6
993.8 | 28.74 | 30.92 | 28.54 | 29.21 | 993.82 | 30.004 | 37.185 | 995.565 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | -0.02 | 1.71
993.81 | 28.69 | 30.86 | 28.41 | 29.33 | 993.79 | 29.903 | 37.311 | 995.537 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.02 | 1.53
993.81 | 28.6 | 30.98 | 28.31 | 29.38 | 993.81 | 29.817 | 37.36 | 995.558 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.29 | 0 | 1.6
993.79 | 28.57 | 31.14 | 28.22 | 29.47 | 993.81 | 29.737 | 37.446 | 995.558 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.35 | -0.02 | 1.67
993.77 | 28.68 | 31.06 | 28.26 | 29.54 | 993.77 | 29.769 | 37.513 | 995.513 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.42 | 0 | 1.52
993.74 | 28.72 | 30.79 | 28.42 | 29.22 | 993.77 | 29.908 | 37.194 | 995.518 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | -0.03 | 1.57
993.72 | 28.55 | 30.92 | 28.26 | 29.27 | 993.73 | 29.775 | 37.252 | 995.474 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.29 | -0.01 | 1.65
993.74 | 28.58 | 31.01 | 28.17 | 29.34 | 993.7 | 29.7 | 37.315 | 995.448 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.41 | 0.04 | 1.67
993.73 | 28.66 | 30.9 | 28.17 | 29.36 | 993.73 | 29.694 | 37.341 | 995.479 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.49 | 0 | 1.54
993.72 | 28.63 | 30.69 | 28.3 | 29.26 | 993.73 | 29.807 | 37.236 | 995.479 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.33 | -0.01 | 1.43
993.72 | 28.44 | 30.67 | 28.14 | 29.22 | 993.73 | 29.673 | 37.193 | 995.473 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | -0.01 | 1.45
993.73 | 28.38 | 30.46 | 27.91 | 29.22 | 993.73 | 29.481 | 37.2 | 995.478 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.47 | 0 | 1.24
993.72 | 28.27 | 30.37 | 28.01 | 28.82 | 993.72 | 29.566 | 36.802 | 995.464 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0 | 1.55
993.72 | 28.15 | 30.33 | 27.83 | 28.85 | 993.73 | 29.411 | 36.832 | 995.476 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.32 | -0.01 | 1.48
993.72 | 28.06 | 30.29 | 27.71 | 28.79 | 993.72 | 29.31 | 36.771 | 995.469 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.35 | 0 | 1.5
993.72 | 27.96 | 30.27 | 27.62 | 28.66 | 993.71 | 29.235 | 36.634 | 995.457 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.34 | 0.01 | 1.61
993.7 | 27.79 | 30.39 | 27.49 | 28.61 | 993.71 | 29.123 | 36.59 | 995.452 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | -0.01 | 1.78
993.7 | 27.72 | 30.37 | 27.42 | 28.63 | 993.69 | 29.07 | 36.603 | 995.434 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0.01 | 1.74
993.69 | 27.72 | 30.3 | 27.25 | 28.62 | 993.7 | 28.92 | 36.601 | 995.448 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.47 | -0.01 | 1.68
993.67 | 27.73 | 30.15 | 27.32 | 28.52 | 993.69 | 28.979 | 36.5 | 995.435 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.41 | -0.02 | 1.63
993.66 | 27.69 | 30.15 | 27.34 | 28.39 | 993.67 | 29 | 36.369 | 995.421 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.35 | -0.01 | 1.76
993.65 | 27.51 | 30.29 | 27.13 | 28.43 | 993.67 | 28.819 | 36.403 | 995.415 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.38 | -0.02 | 1.86
993.65 | 27.54 | 30.3 | 27.02 | 28.54 | 993.65 | 28.733 | 36.519 | 995.401 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.52 | 0 | 1.76
993.65 | 27.51 | 30.3 | 27.11 | 28.38 | 993.67 | 28.803 | 36.357 | 995.42 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.4 | -0.02 | 1.92
993.63 | 27.44 | 30.33 | 26.97 | 28.52 | 993.65 | 28.69 | 36.496 | 995.394 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.47 | -0.02 | 1.81
993.63 | 27.44 | 30.27 | 26.86 | 28.74 | 993.64 | 28.6 | 36.721 | 995.383 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.58 | -0.01 | 1.53
993.62 | 27.45 | 30.2 | 26.97 | 28.45 | 993.63 | 28.685 | 36.43 | 995.376 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.48 | -0.01 | 1.75
993.62 | 27.4 | 30.14 | 26.95 | 28.35 | 993.62 | 28.669 | 36.327 | 995.371 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.45 | 0 | 1.79
993.63 | 27.35 | 30.16 | 26.81 | 28.48 | 993.61 | 28.552 | 36.462 | 995.36 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.54 | 0.02 | 1.68
993.63 | 27.36 | 30.15 | 26.81 | 28.47 | 993.62 | 28.557 | 36.451 | 995.367 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.55 | 0.01 | 1.68
993.63 | 27.3 | 30.12 | 26.8 | 28.43 | 993.62 | 28.541 | 36.409 | 995.37 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.5 | 0.01 | 1.69
993.61 | 27.27 | 30.17 | 26.68 | 28.39 | 993.63 | 28.445 | 36.37 | 995.372 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.59 | -0.02 | 1.78
993.62 | 27.31 | 30.09 | 26.6 | 28.45 | 993.61 | 28.381 | 36.43 | 995.352 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.71 | 0.01 | 1.64
993.62 | 27.25 | 30.06 | 26.73 | 28.35 | 993.6 | 28.487 | 36.323 | 995.348 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.52 | 0.02 | 1.71
993.6 | 27.15 | 30.13 | 26.5 | 28.54 | 993.58 | 28.295 | 36.521 | 995.331 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.65 | 0.02 | 1.59
993.59 | 27.15 | 30.11 | 26.55 | 28.28 | 993.57 | 28.333 | 36.26 | 995.316 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.6 | 0.02 | 1.83
993.58 | 27.08 | 30.16 | 26.56 | 28.37 | 993.55 | 28.343 | 36.351 | 995.298 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.52 | 0.03 | 1.79
993.56 | 27.16 | 30.04 | 26.56 | 28.33 | 993.54 | 28.343 | 36.308 | 995.291 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.6 | 0.02 | 1.71
993.55 | 27.15 | 29.94 | 26.56 | 28.18 | 993.53 | 28.343 | 36.155 | 995.277 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.59 | 0.02 | 1.76
993.54 | 27.07 | 29.99 | 26.55 | 28.23 | 993.53 | 28.333 | 36.205 | 995.276 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.52 | 0.01 | 1.76
993.53 | 27.02 | 30.02 | 26.41 | 28.32 | 993.51 | 28.22 | 36.292 | 995.257 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.61 | 0.02 | 1.7
993.52 | 27.01 | 30.05 | 26.38 | 28.37 | 993.5 | 28.188 | 36.348 | 995.248 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.63 | 0.02 | 1.68
993.52 | 27.01 | 29.99 | 26.46 | 28.22 | 993.5 | 28.263 | 36.195 | 995.244 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.55 | 0.02 | 1.77
993.52 | 26.96 | 30.07 | 26.36 | 28.23 | 993.49 | 28.178 | 36.206 | 995.235 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.6 | 0.03 | 1.84
993.49 | 26.94 | 30.12 | 26.43 | 28.23 | 993.47 | 28.236 | 36.203 | 995.214 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.51 | 0.02 | 1.89
993.48 | 26.98 | 29.99 | 26.45 | 28.19 | 993.46 | 28.252 | 36.172 | 995.202 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.53 | 0.02 | 1.8
993.46 | 26.92 | 30.05 | 26.34 | 28.08 | 993.45 | 28.156 | 36.06 | 995.198 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.58 | 0.01 | 1.97
993.45 | 26.96 | 30.01 | 26.17 | 28.43 | 993.42 | 28.012 | 36.407 | 995.166 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.79 | 0.03 | 1.58
993.45 | 27 | 29.77 | 26.35 | 28.15 | 993.42 | 28.167 | 36.132 | 995.162 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.65 | 0.03 | 1.62
993.43 | 26.91 | 29.32 | 26.28 | 27.99 | 993.41 | 28.108 | 35.97 | 995.157 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.63 | 0.02 | 1.33
993.43 | 26.61 | 29.19 | 26.04 | 27.74 | 993.4 | 27.905 | 35.72 | 995.146 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.57 | 0.03 | 1.45
993.43 | 26.47 | 29.39 | 25.84 | 27.74 | 993.38 | 27.734 | 35.72 | 995.128 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.63 | 0.05 | 1.65
993.43 | 26.48 | 29.48 | 25.79 | 27.63 | 993.39 | 27.697 | 35.611 | 995.139 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.69 | 0.04 | 1.85
993.43 | 26.52 | 29.66 | 25.75 | 27.66 | 993.39 | 27.665 | 35.638 | 995.139 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.77 | 0.04 | 2
993.43 | 26.57 | 29.61 | 25.96 | 27.7 | 993.39 | 27.836 | 35.672 | 995.135 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.61 | 0.04 | 1.91
993.44 | 26.57 | 29.71 | 26.03 | 27.63 | 993.38 | 27.895 | 35.609 | 995.125 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.54 | 0.06 | 2.08
993.43 | 26.65 | 29.7 | 25.89 | 27.82 | 993.41 | 27.782 | 35.8 | 995.152 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.76 | 0.02 | 1.88
993.44 | 26.74 | 29.59 | 26.06 | 27.67 | 993.38 | 27.927 | 35.651 | 995.129 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.68 | 0.06 | 1.92
993.43 | 26.74 | 29.57 | 26.03 | 27.79 | 993.4 | 27.9 | 35.77 | 995.149 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.71 | 0.03 | 1.78
993.44 | 26.64 | 29.7 | 25.99 | 27.81 | 993.39 | 27.863 | 35.791 | 995.139 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.65 | 0.05 | 1.89
993.45 | 26.54 | 29.87 | 25.75 | 28.09 | 993.4 | 27.665 | 36.063 | 995.147 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.79 | 0.05 | 1.78
993.44 | 26.58 | 29.89 | 25.77 | 28.1 | 993.42 | 27.681 | 36.081 | 995.167 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.81 | 0.02 | 1.79
993.45 | 26.65 | 29.8 | 25.87 | 28.02 | 993.41 | 27.761 | 35.997 | 995.154 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.78 | 0.04 | 1.78
993.46 | 26.56 | 29.85 | 25.84 | 28.07 | 993.4 | 27.74 | 36.046 | 995.149 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.72 | 0.06 | 1.78
993.46 | 26.47 | 29.98 | 25.72 | 28.17 | 993.41 | 27.638 | 36.145 | 995.161 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.75 | 0.05 | 1.81
993.44 | 26.57 | 29.85 | 25.64 | 28.29 | 993.42 | 27.569 | 36.269 | 995.162 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.93 | 0.02 | 1.56
993.45 | 26.51 | 29.84 | 25.81 | 28.01 | 993.39 | 27.713 | 35.985 | 995.14 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.7 | 0.06 | 1.83
993.46 | 26.47 | 29.83 | 25.75 | 28.04 | 993.4 | 27.66 | 36.021 | 995.147 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.72 | 0.06 | 1.79
993.45 | 26.51 | 29.8 | 25.77 | 28.08 | 993.39 | 27.676 | 36.056 | 995.14 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.74 | 0.06 | 1.72
993.44 | 26.44 | 29.8 | 25.72 | 28.02 | 993.4 | 27.638 | 35.995 | 995.148 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.72 | 0.04 | 1.78
993.44 | 26.34 | 29.94 | 25.67 | 28.02 | 993.42 | 27.596 | 35.993 | 995.162 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.67 | 0.02 | 1.92
993.43 | 26.35 | 29.94 | 25.56 | 28.11 | 993.41 | 27.505 | 36.086 | 995.154 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.79 | 0.02 | 1.83
993.43 | 26.44 | 29.83 | 25.56 | 28.12 | 993.4 | 27.505 | 36.094 | 995.148 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.88 | 0.03 | 1.71
993.41 | 26.41 | 29.81 | 25.63 | 27.95 | 993.39 | 27.564 | 35.93 | 995.135 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.78 | 0.02 | 1.86
993.41 | 26.41 | 29.79 | 25.54 | 28.04 | 993.37 | 27.483 | 36.018 | 995.121 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.87 | 0.04 | 1.75
993.39 | 26.36 | 29.77 | 25.59 | 27.92 | 993.37 | 27.526 | 35.899 | 995.112 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.77 | 0.02 | 1.85
993.4 | 26.34 | 29.79 | 25.59 | 27.98 | 993.36 | 27.531 | 35.958 | 995.109 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.75 | 0.04 | 1.81
993.4 | 26.27 | 29.87 | 25.52 | 27.98 | 993.35 | 27.467 | 35.953 | 995.1 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.75 | 0.05 | 1.89
993.39 | 26.33 | 29.85 | 25.51 | 28.01 | 993.36 | 27.462 | 35.987 | 995.102 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.82 | 0.03 | 1.84
993.39 | 26.28 | 29.86 | 25.52 | 27.98 | 993.36 | 27.473 | 35.953 | 995.102 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.76 | 0.03 | 1.88
993.39 | 26.2 | 29.95 | 25.32 | 28.17 | 993.35 | 27.302 | 36.147 | 995.092 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.88 | 0.04 | 1.78
993.4 | 26.18 | 30.1 | 25.35 | 28.14 | 993.32 | 27.323 | 36.119 | 995.064 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.83 | 0.08 | 1.96
993.43 | 26.23 | 30.09 | 25.36 | 28.27 | 993.34 | 27.334 | 36.248 | 995.084 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.87 | 0.09 | 1.82
993.4 | 26.32 | 30.04 | 25.41 | 28.19 | 993.34 | 27.377 | 36.17 | 995.091 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.91 | 0.06 | 1.85
993.38 | 26.29 | 30.1 | 25.52 | 28.19 | 993.35 | 27.473 | 36.166 | 995.095 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.77 | 0.03 | 1.91
993.39 | 26.3 | 30.07 | 25.49 | 28.27 | 993.32 | 27.441 | 36.248 | 995.071 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.81 | 0.07 | 1.8
993.38 | 26.27 | 30.03 | 25.51 | 28.15 | 993.34 | 27.462 | 36.13 | 995.085 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.76 | 0.04 | 1.88
993.38 | 26.25 | 29.99 | 25.5 | 28.18 | 993.33 | 27.451 | 36.159 | 995.077 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.75 | 0.05 | 1.81
993.4 | 26.26 | 29.96 | 25.42 | 28.18 | 993.33 | 27.387 | 36.155 | 995.075 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.84 | 0.07 | 1.78
993.41 | 26.14 | 30.02 | 25.28 | 28.28 | 993.36 | 27.264 | 36.252 | 995.103 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.86 | 0.05 | 1.74
993.4 | 26.13 | 30.06 | 25.28 | 28.33 | 993.38 | 27.264 | 36.311 | 995.126 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.85 | 0.02 | 1.73
993.42 | 26.17 | 29.96 | 25.21 | 28.24 | 993.37 | 27.211 | 36.22 | 995.113 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.96 | 0.05 | 1.72
993.41 | 26.13 | 30.02 | 25.24 | 28.21 | 993.39 | 27.232 | 36.191 | 995.137 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.89 | 0.02 | 1.81
993.41 | 26.08 | 30.08 | 25.21 | 28.23 | 993.39 | 27.211 | 36.206 | 995.139 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.87 | 0.02 | 1.85
993.42 | 26.07 | 30.02 | 25.18 | 28.33 | 993.39 | 27.179 | 36.308 | 995.141 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.89 | 0.03 | 1.69
993.43 | 26.03 | 30.12 | 25.16 | 28.28 | 993.38 | 27.163 | 36.26 | 995.129 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.87 | 0.05 | 1.84
993.43 | 26.08 | 30.04 | 25.21 | 28.32 | 993.39 | 27.211 | 36.294 | 995.138 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.87 | 0.04 | 1.72
993.43 | 26.11 | 29.92 | 25.19 | 28.19 | 993.4 | 27.195 | 36.17 | 995.145 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.92 | 0.03 | 1.73
993.44 | 26.14 | 29.87 | 25.28 | 28 | 993.4 | 27.264 | 35.981 | 995.145 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.86 | 0.04 | 1.87
993.43 | 26.1 | 29.81 | 25.23 | 28.04 | 993.42 | 27.227 | 36.018 | 995.163 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.87 | 0.01 | 1.77
993.43 | 26.06 | 29.85 | 25.05 | 28.23 | 993.38 | 27.078 | 36.212 | 995.126 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.01 | 0.05 | 1.62
993.44 | 25.99 | 29.94 | 25.15 | 28.06 | 993.4 | 27.158 | 36.039 | 995.151 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.84 | 0.04 | 1.88
993.43 | 26 | 29.95 | 24.94 | 28.29 | 993.4 | 26.981 | 36.267 | 995.146 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.06 | 0.03 | 1.66
993.44 | 26.18 | 29.88 | 25.15 | 28.03 | 993.4 | 27.158 | 36.012 | 995.148 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.03 | 0.04 | 1.85
993.44 | 26.71 | 29.28 | 25.51 | 27.68 | 993.38 | 27.462 | 35.657 | 995.125 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.2 | 0.06 | 1.6
993.44 | 27.3 | 28.51 | 25.94 | 27.32 | 993.39 | 27.825 | 35.295 | 995.141 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.36 | 0.05 | 1.19
993.46 | 27.8 | 27.76 | 26.38 | 26.8 | 993.38 | 28.194 | 34.774 | 995.131 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.42 | 0.08 | 0.96
993.46 | 28.28 | 27.1 | 26.81 | 26.3 | 993.42 | 28.552 | 34.282 | 995.166 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.47 | 0.04 | 0.8
993.46 | 28.68 | 26.52 | 27.2 | 25.88 | 993.43 | 28.877 | 33.857 | 995.175 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.48 | 0.03 | 0.64
993.46 | 29.04 | 26.01 | 27.56 | 25.48 | 993.42 | 29.182 | 33.454 | 995.167 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.48 | 0.04 | 0.53
993.47 | 29.39 | 25.6 | 27.89 | 25.11 | 993.43 | 29.465 | 33.084 | 995.181 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.5 | 0.04 | 0.49
993.46 | 29.69 | 25.25 | 28.23 | 24.89 | 993.41 | 29.748 | 32.863 | 995.159 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.46 | 0.05 | 0.36
993.45 | 29.95 | 24.97 | 28.5 | 24.72 | 993.4 | 29.972 | 32.695 | 995.148 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.45 | 0.05 | 0.25
993.47 | 30.18 | 24.72 | 28.74 | 24.55 | 993.42 | 30.18 | 32.525 | 995.163 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.44 | 0.05 | 0.17
993.45 | 30.38 | 24.46 | 29 | 24.32 | 993.43 | 30.394 | 32.294 | 995.172 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.38 | 0.02 | 0.14
993.46 | 30.59 | 24.3 | 29.23 | 24.18 | 993.39 | 30.592 | 32.161 | 995.139 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.36 | 0.07 | 0.12
993.46 | 30.75 | 24.1 | 29.42 | 23.96 | 993.4 | 30.747 | 31.942 | 995.15 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.33 | 0.06 | 0.14
993.46 | 30.92 | 23.97 | 29.62 | 23.87 | 993.39 | 30.912 | 31.846 | 995.141 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.3 | 0.07 | 0.1
993.47 | 31.07 | 23.79 | 29.78 | 23.8 | 993.4 | 31.051 | 31.776 | 995.146 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.29 | 0.07 | -0.01
993.48 | 31.2 | 23.66 | 29.95 | 23.68 | 993.4 | 31.19 | 31.654 | 995.143 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.25 | 0.08 | -0.02
993.45 | 31.32 | 23.61 | 30.1 | 23.65 | 993.41 | 31.323 | 31.623 | 995.155 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.22 | 0.04 | -0.04
993.46 | 31.43 | 23.55 | 30.26 | 23.53 | 993.4 | 31.457 | 31.51 | 995.144 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.17 | 0.06 | 0.02
993.47 | 31.55 | 23.52 | 30.34 | 23.53 | 993.4 | 31.521 | 31.503 | 995.142 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.21 | 0.07 | -0.01
993.46 | 31.64 | 23.47 | 30.49 | 23.59 | 993.42 | 31.644 | 31.566 | 995.164 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.15 | 0.04 | -0.12
993.47 | 31.74 | 23.43 | 30.57 | 23.54 | 993.41 | 31.713 | 31.512 | 995.157 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.17 | 0.06 | -0.11
993.48 | 31.85 | 23.43 | 30.7 | 23.47 | 993.39 | 31.825 | 31.446 | 995.138 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.15 | 0.09 | -0.04
993.48 | 31.91 | 23.37 | 30.82 | 23.42 | 993.41 | 31.927 | 31.4 | 995.16 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.09 | 0.07 | -0.05
993.47 | 31.99 | 23.34 | 30.9 | 23.39 | 993.41 | 31.991 | 31.367 | 995.156 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.09 | 0.06 | -0.05
993.48 | 32.06 | 23.34 | 30.98 | 23.43 | 993.39 | 32.06 | 31.409 | 995.138 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.08 | 0.09 | -0.09
993.47 | 32.11 | 23.27 | 31.09 | 23.37 | 993.4 | 32.151 | 31.352 | 995.145 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.02 | 0.07 | -0.1
993.45 | 32.2 | 23.24 | 31.19 | 23.29 | 993.4 | 32.237 | 31.27 | 995.151 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1.01 | 0.05 | -0.05
993.45 | 32.27 | 23.22 | 31.29 | 23.26 | 993.39 | 32.317 | 31.242 | 995.138 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.98 | 0.06 | -0.04
993.46 | 32.35 | 23.18 | 31.35 | 23.27 | 993.38 | 32.37 | 31.251 | 995.123 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 1 | 0.08 | -0.09
993.46 | 32.41 | 23.17 | 31.42 | 23.31 | 993.37 | 32.429 | 31.285 | 995.116 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.99 | 0.09 | -0.14
993.47 | 32.44 | 23.08 | 31.51 | 23.25 | 993.38 | 32.504 | 31.232 | 995.131 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.93 | 0.09 | -0.17
993.46 | 32.52 | 23.16 | 31.59 | 23.18 | 993.41 | 32.573 | 31.16 | 995.159 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.93 | 0.05 | -0.02
993.46 | 32.57 | 23.1 | 31.68 | 23.2 | 993.41 | 32.648 | 31.179 | 995.153 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.89 | 0.05 | -0.1
993.46 | 32.62 | 23.06 | 31.72 | 23.12 | 993.4 | 32.68 | 31.1 | 995.144 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.9 | 0.06 | -0.06
993.45 | 32.66 | 23.05 | 31.82 | 23.08 | 993.39 | 32.765 | 31.053 | 995.14 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.84 | 0.06 | -0.03
993.47 | 32.73 | 23.05 | 31.84 | 23.13 | 993.39 | 32.781 | 31.112 | 995.141 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.89 | 0.08 | -0.08
993.46 | 32.75 | 23.06 | 31.93 | 23.12 | 993.42 | 32.856 | 31.093 | 995.162 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.82 | 0.04 | -0.06
993.46 | 32.8 | 23.04 | 32 | 23.14 | 993.41 | 32.92 | 31.119 | 995.156 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.8 | 0.05 | -0.1
993.47 | 32.83 | 23.04 | 32.02 | 23.07 | 993.42 | 32.931 | 31.047 | 995.166 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.81 | 0.05 | -0.03
993.46 | 32.89 | 23.02 | 32.09 | 23.1 | 993.41 | 32.99 | 31.076 | 995.152 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.8 | 0.05 | -0.08
993.46 | 32.93 | 23.04 | 32.11 | 23.07 | 993.4 | 33.011 | 31.047 | 995.15 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.82 | 0.06 | -0.03
993.46 | 32.97 | 23.03 | 32.21 | 23.05 | 993.41 | 33.096 | 31.03 | 995.157 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.76 | 0.05 | -0.02
993.46 | 33 | 23.04 | 32.24 | 23.09 | 993.4 | 33.118 | 31.062 | 995.147 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.76 | 0.06 | -0.05
993.46 | 33.05 | 23.04 | 32.28 | 23.06 | 993.38 | 33.155 | 31.041 | 995.13 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.77 | 0.08 | -0.02
993.46 | 33.07 | 23 | 32.35 | 23.07 | 993.38 | 33.209 | 31.045 | 995.122 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.72 | 0.08 | -0.07
993.47 | 33.12 | 23.01 | 32.36 | 23.04 | 993.37 | 33.219 | 31.015 | 995.115 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.76 | 0.1 | -0.03
993.46 | 33.15 | 23.02 | 32.44 | 23.09 | 993.39 | 33.289 | 31.07 | 995.141 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.71 | 0.07 | -0.07
993.48 | 33.19 | 23.04 | 32.47 | 23.07 | 993.4 | 33.315 | 31.047 | 995.148 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.72 | 0.08 | -0.03
993.49 | 33.22 | 23.04 | 32.49 | 23.11 | 993.42 | 33.331 | 31.091 | 995.162 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.73 | 0.07 | -0.07
993.5 | 33.26 | 22.97 | 32.59 | 23.04 | 993.44 | 33.417 | 31.02 | 995.183 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.67 | 0.06 | -0.07
993.52 | 33.29 | 23.04 | 32.59 | 23.04 | 993.44 | 33.417 | 31.018 | 995.191 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.7 | 0.08 | 0
993.52 | 33.32 | 23.03 | 32.64 | 23.06 | 993.46 | 33.46 | 31.036 | 995.207 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.68 | 0.06 | -0.03
993.53 | 33.35 | 23.04 | 32.65 | 23.07 | 993.46 | 33.465 | 31.047 | 995.203 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.7 | 0.07 | -0.03
993.56 | 33.38 | 23.06 | 32.73 | 23.08 | 993.48 | 33.529 | 31.058 | 995.222 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.65 | 0.08 | -0.02
993.57 | 33.4 | 23.06 | 32.73 | 23.11 | 993.52 | 33.529 | 31.089 | 995.262 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.67 | 0.05 | -0.05
993.58 | 33.42 | 23.07 | 32.8 | 23.13 | 993.52 | 33.588 | 31.104 | 995.262 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.62 | 0.06 | -0.06
993.6 | 33.46 | 23.08 | 32.84 | 23.06 | 993.54 | 33.625 | 31.037 | 995.288 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.62 | 0.06 | 0.02
993.61 | 33.47 | 23.07 | 32.85 | 23.13 | 993.55 | 33.636 | 31.106 | 995.294 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.62 | 0.06 | -0.06
993.61 | 33.51 | 23.08 | 32.91 | 23.08 | 993.57 | 33.679 | 31.057 | 995.315 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.6 | 0.04 | 0
993.63 | 33.56 | 23.11 | 32.92 | 23.09 | 993.56 | 33.689 | 31.064 | 995.306 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.64 | 0.07 | 0.02
993.65 | 33.58 | 23.1 | 32.99 | 23.13 | 993.59 | 33.753 | 31.106 | 995.332 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.59 | 0.06 | -0.03
993.66 | 33.58 | 23.1 | 33.02 | 23.13 | 993.62 | 33.775 | 31.108 | 995.365 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.56 | 0.04 | -0.03
993.67 | 33.62 | 23.09 | 33.05 | 23.16 | 993.64 | 33.801 | 31.139 | 995.387 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.57 | 0.03 | -0.07
993.68 | 33.64 | 23.11 | 33.05 | 23.15 | 993.63 | 33.801 | 31.131 | 995.38 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.59 | 0.05 | -0.04
993.69 | 33.67 | 23.12 | 33.07 | 23.21 | 993.66 | 33.817 | 31.188 | 995.402 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.6 | 0.03 | -0.09
993.7 | 33.66 | 23.13 | 33.11 | 23.21 | 993.65 | 33.849 | 31.182 | 995.397 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.55 | 0.05 | -0.08
993.7 | 33.68 | 23.13 | 33.15 | 23.19 | 993.65 | 33.882 | 31.163 | 995.397 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.53 | 0.05 | -0.06
993.72 | 33.75 | 23.17 | 33.18 | 23.14 | 993.65 | 33.908 | 31.118 | 995.401 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.57 | 0.07 | 0.03
993.72 | 33.74 | 23.15 | 33.25 | 23.19 | 993.67 | 33.967 | 31.163 | 995.413 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.49 | 0.05 | -0.04
993.71 | 33.78 | 23.15 | 33.23 | 23.2 | 993.67 | 33.951 | 31.177 | 995.414 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.55 | 0.04 | -0.05
993.72 | 33.81 | 23.15 | 33.25 | 23.16 | 993.66 | 33.967 | 31.133 | 995.406 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.56 | 0.06 | -0.01
993.73 | 33.83 | 23.12 | 33.32 | 23.15 | 993.66 | 34.031 | 31.127 | 995.411 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.51 | 0.07 | -0.03
993.73 | 33.85 | 23.16 | 33.3 | 23.17 | 993.67 | 34.01 | 31.142 | 995.421 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.55 | 0.06 | -0.01
993.74 | 33.84 | 23.18 | 33.34 | 23.19 | 993.67 | 34.047 | 31.169 | 995.418 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.5 | 0.07 | -0.01
993.75 | 33.86 | 23.2 | 33.34 | 23.17 | 993.7 | 34.042 | 31.144 | 995.446 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.52 | 0.05 | 0.03
993.75 | 33.87 | 23.16 | 33.39 | 23.19 | 993.69 | 34.09 | 31.169 | 995.44 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.48 | 0.06 | -0.03
993.77 | 33.92 | 23.2 | 33.39 | 23.22 | 993.7 | 34.09 | 31.194 | 995.444 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.53 | 0.07 | -0.02
993.79 | 33.94 | 23.19 | 33.46 | 23.2 | 993.72 | 34.143 | 31.181 | 995.47 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.48 | 0.07 | -0.01
993.81 | 33.93 | 23.23 | 33.46 | 23.17 | 993.74 | 34.143 | 31.152 | 995.49 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.47 | 0.07 | 0.06
993.82 | 33.98 | 23.21 | 33.49 | 23.26 | 993.74 | 34.17 | 31.238 | 995.491 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.49 | 0.08 | -0.05
993.83 | 33.97 | 23.24 | 33.53 | 23.25 | 993.76 | 34.207 | 31.226 | 995.503 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.44 | 0.07 | -0.01
993.85 | 33.98 | 23.23 | 33.5 | 23.22 | 993.77 | 34.181 | 31.198 | 995.512 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.48 | 0.08 | 0.01
993.85 | 33.99 | 23.25 | 33.55 | 23.28 | 993.78 | 34.218 | 31.257 | 995.523 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.44 | 0.07 | -0.03
993.86 | 34.02 | 23.25 | 33.54 | 23.33 | 993.8 | 34.213 | 31.304 | 995.547 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.48 | 0.06 | -0.08
993.88 | 34.05 | 23.27 | 33.59 | 23.24 | 993.8 | 34.255 | 31.213 | 995.549 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.46 | 0.08 | 0.03
993.89 | 34.08 | 23.25 | 33.65 | 23.26 | 993.83 | 34.303 | 31.236 | 995.579 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.43 | 0.06 | -0.01
993.9 | 34.09 | 23.25 | 33.65 | 23.24 | 993.83 | 34.309 | 31.217 | 995.579 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.44 | 0.07 | 0.01
993.91 | 34.12 | 23.23 | 33.67 | 23.29 | 993.84 | 34.319 | 31.264 | 995.588 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.45 | 0.07 | -0.06
993.92 | 34.11 | 23.28 | 33.69 | 23.28 | 993.85 | 34.335 | 31.255 | 995.601 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.42 | 0.07 | 0
993.94 | 34.13 | 23.32 | 33.67 | 23.28 | 993.88 | 34.325 | 31.261 | 995.631 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.46 | 0.06 | 0.04
993.96 | 34.14 | 23.3 | 33.72 | 23.33 | 993.9 | 34.362 | 31.308 | 995.645 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.42 | 0.06 | -0.03
993.97 | 34.14 | 23.31 | 33.75 | 23.31 | 993.89 | 34.389 | 31.284 | 995.633 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.39 | 0.08 | 0
993.98 | 34.14 | 23.33 | 33.74 | 23.32 | 993.91 | 34.384 | 31.299 | 995.657 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.4 | 0.07 | 0.01
994 | 34.19 | 23.31 | 33.77 | 23.33 | 993.95 | 34.405 | 31.312 | 995.696 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.42 | 0.05 | -0.02
994.01 | 34.2 | 23.34 | 33.76 | 23.35 | 993.94 | 34.4 | 31.325 | 995.691 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.44 | 0.07 | -0.01
994.02 | 34.2 | 23.35 | 33.79 | 23.35 | 993.96 | 34.421 | 31.327 | 995.708 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.41 | 0.06 | 0
994.04 | 34.23 | 23.37 | 33.81 | 23.35 | 993.98 | 34.437 | 31.324 | 995.728 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.42 | 0.06 | 0.02
994.05 | 34.24 | 23.34 | 33.81 | 23.37 | 994 | 34.442 | 31.346 | 995.746 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.43 | 0.05 | -0.03
994.04 | 34.26 | 23.39 | 33.86 | 23.38 | 993.98 | 34.48 | 31.356 | 995.73 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.4 | 0.06 | 0.01
994.05 | 34.26 | 23.41 | 33.86 | 23.35 | 993.99 | 34.48 | 31.325 | 995.735 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.4 | 0.06 | 0.06
994.07 | 34.27 | 23.37 | 33.88 | 23.41 | 994 | 34.496 | 31.385 | 995.745 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.39 | 0.07 | -0.04
994.07 | 34.28 | 23.39 | 33.87 | 23.38 | 994.01 | 34.49 | 31.362 | 995.754 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.41 | 0.06 | 0.01
994.09 | 34.29 | 23.44 | 33.89 | 23.38 | 994 | 34.506 | 31.356 | 995.751 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.4 | 0.09 | 0.06
994.1 | 34.31 | 23.45 | 33.94 | 23.46 | 994.03 | 34.549 | 31.432 | 995.777 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.37 | 0.07 | -0.01
994.11 | 34.32 | 23.43 | 33.93 | 23.43 | 994.04 | 34.544 | 31.411 | 995.782 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.39 | 0.07 | 0
994.12 | 34.34 | 23.45 | 33.98 | 23.42 | 994.05 | 34.581 | 31.394 | 995.793 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.36 | 0.07 | 0.03
994.13 | 34.34 | 23.44 | 33.96 | 23.41 | 994.07 | 34.565 | 31.39 | 995.812 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.38 | 0.06 | 0.03
994.13 | 34.36 | 23.45 | 33.97 | 23.4 | 994.08 | 34.57 | 31.375 | 995.828 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.39 | 0.05 | 0.05
994.14 | 34.37 | 23.46 | 34.01 | 23.47 | 994.08 | 34.608 | 31.448 | 995.829 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.36 | 0.06 | -0.01
994.15 | 34.38 | 23.52 | 33.99 | 23.47 | 994.08 | 34.592 | 31.451 | 995.829 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.39 | 0.07 | 0.05
994.16 | 34.39 | 23.53 | 34.01 | 23.5 | 994.09 | 34.608 | 31.478 | 995.84 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.38 | 0.07 | 0.03
994.17 | 34.41 | 23.51 | 34.04 | 23.52 | 994.1 | 34.635 | 31.495 | 995.85 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.37 | 0.07 | -0.01
994.19 | 34.41 | 23.51 | 34.04 | 23.49 | 994.12 | 34.635 | 31.467 | 995.871 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.37 | 0.07 | 0.02
994.21 | 34.42 | 23.51 | 34.04 | 23.54 | 994.14 | 34.635 | 31.514 | 995.888 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.38 | 0.07 | -0.03
994.23 | 34.44 | 23.52 | 34.11 | 23.52 | 994.16 | 34.688 | 31.499 | 995.903 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.33 | 0.07 | 0
994.24 | 34.47 | 23.54 | 34.09 | 23.54 | 994.19 | 34.672 | 31.518 | 995.933 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.38 | 0.05 | 0
994.25 | 34.45 | 23.54 | 34.11 | 23.51 | 994.18 | 34.688 | 31.491 | 995.931 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.34 | 0.07 | 0.03
994.26 | 34.47 | 23.54 | 34.13 | 23.47 | 994.2 | 34.709 | 31.449 | 995.946 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.34 | 0.06 | 0.07
994.27 | 34.49 | 23.55 | 34.17 | 23.51 | 994.19 | 34.741 | 31.486 | 995.936 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.32 | 0.08 | 0.04
994.27 | 34.48 | 23.55 | 34.15 | 23.54 | 994.21 | 34.725 | 31.516 | 995.954 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.33 | 0.06 | 0.01
994.27 | 34.51 | 23.58 | 34.14 | 23.58 | 994.2 | 34.715 | 31.552 | 995.95 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.37 | 0.07 | 0
994.27 | 34.5 | 23.6 | 34.16 | 23.54 | 994.2 | 34.736 | 31.522 | 995.949 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.34 | 0.07 | 0.06
994.27 | 34.53 | 23.62 | 34.15 | 23.61 | 994.21 | 34.725 | 31.583 | 995.956 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.38 | 0.06 | 0.01
994.28 | 34.52 | 23.63 | 34.18 | 23.61 | 994.22 | 34.752 | 31.589 | 995.968 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.34 | 0.06 | 0.02
994.28 | 34.52 | 23.64 | 34.21 | 23.64 | 994.22 | 34.779 | 31.621 | 995.97 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.31 | 0.06 | 0
994.3 | 34.52 | 23.67 | 34.2 | 23.64 | 994.24 | 34.768 | 31.619 | 995.983 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.32 | 0.06 | 0.03
994.32 | 34.53 | 23.68 | 34.19 | 23.72 | 994.26 | 34.763 | 31.697 | 996.002 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.34 | 0.06 | -0.04
994.33 | 34.56 | 23.65 | 34.23 | 23.63 | 994.27 | 34.789 | 31.604 | 996.017 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.33 | 0.06 | 0.02
994.35 | 34.55 | 23.65 | 34.25 | 23.67 | 994.28 | 34.811 | 31.646 | 996.028 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0.07 | -0.02
994.36 | 34.58 | 23.65 | 34.24 | 23.66 | 994.3 | 34.8 | 31.638 | 996.042 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.34 | 0.06 | -0.01
994.37 | 34.58 | 23.68 | 34.26 | 23.69 | 994.31 | 34.816 | 31.665 | 996.055 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.32 | 0.06 | -0.01
994.38 | 34.58 | 23.69 | 34.25 | 23.68 | 994.32 | 34.805 | 31.659 | 996.065 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.33 | 0.06 | 0.01
994.38 | 34.59 | 23.71 | 34.25 | 23.71 | 994.33 | 34.811 | 31.69 | 996.073 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.34 | 0.05 | 0
994.4 | 34.61 | 23.72 | 34.26 | 23.7 | 994.34 | 34.816 | 31.678 | 996.083 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.35 | 0.06 | 0.02
994.4 | 34.6 | 23.72 | 34.3 | 23.74 | 994.35 | 34.854 | 31.713 | 996.099 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0.05 | -0.02
994.39 | 34.6 | 23.71 | 34.32 | 23.7 | 994.32 | 34.864 | 31.682 | 996.07 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.07 | 0.01
994.39 | 34.61 | 23.77 | 34.32 | 23.72 | 994.35 | 34.864 | 31.694 | 996.094 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.29 | 0.04 | 0.05
994.42 | 34.62 | 23.81 | 34.3 | 23.78 | 994.33 | 34.854 | 31.757 | 996.079 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.32 | 0.09 | 0.03
994.43 | 34.63 | 23.76 | 34.32 | 23.73 | 994.36 | 34.87 | 31.709 | 996.106 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.31 | 0.07 | 0.03
994.43 | 34.65 | 23.76 | 34.35 | 23.74 | 994.38 | 34.896 | 31.718 | 996.13 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0.05 | 0.02
994.42 | 34.64 | 23.8 | 34.34 | 23.76 | 994.39 | 34.886 | 31.736 | 996.135 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0.03 | 0.04
994.43 | 34.63 | 23.83 | 34.33 | 23.79 | 994.36 | 34.875 | 31.762 | 996.111 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0.07 | 0.04
994.45 | 34.65 | 23.84 | 34.33 | 23.84 | 994.38 | 34.88 | 31.818 | 996.126 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.32 | 0.07 | 0
994.46 | 34.66 | 23.83 | 34.36 | 23.82 | 994.38 | 34.902 | 31.8 | 996.13 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0.08 | 0.01
994.45 | 34.66 | 23.83 | 34.4 | 23.82 | 994.4 | 34.934 | 31.793 | 996.148 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0.05 | 0.01
994.47 | 34.69 | 23.84 | 34.38 | 23.79 | 994.4 | 34.918 | 31.768 | 996.146 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.31 | 0.07 | 0.05
994.46 | 34.69 | 23.87 | 34.41 | 23.84 | 994.42 | 34.944 | 31.816 | 996.164 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.04 | 0.03
994.5 | 34.69 | 23.9 | 34.42 | 23.83 | 994.4 | 34.95 | 31.812 | 996.151 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | 0.1 | 0.07
994.48 | 34.71 | 23.89 | 34.4 | 23.87 | 994.43 | 34.939 | 31.844 | 996.18 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.31 | 0.05 | 0.02
994.49 | 34.72 | 23.9 | 34.4 | 23.89 | 994.43 | 34.939 | 31.867 | 996.179 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.32 | 0.06 | 0.01
994.47 | 34.7 | 23.95 | 34.42 | 23.86 | 994.43 | 34.95 | 31.839 | 996.177 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.04 | 0.09
994.49 | 34.72 | 23.95 | 34.42 | 23.97 | 994.4 | 34.95 | 31.947 | 996.151 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0.09 | -0.02
994.49 | 34.73 | 23.94 | 34.45 | 23.88 | 994.43 | 34.982 | 31.854 | 996.175 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.06 | 0.06
994.52 | 34.73 | 23.94 | 34.45 | 23.91 | 994.42 | 34.976 | 31.888 | 996.162 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.1 | 0.03
994.52 | 34.76 | 23.91 | 34.47 | 23.93 | 994.45 | 34.998 | 31.911 | 996.192 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.29 | 0.07 | -0.02
994.51 | 34.75 | 23.96 | 34.49 | 23.91 | 994.46 | 35.008 | 31.886 | 996.209 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0.05 | 0.05
994.55 | 34.74 | 23.95 | 34.47 | 23.96 | 994.44 | 34.998 | 31.94 | 996.183 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | 0.11 | -0.01
994.52 | 34.77 | 23.97 | 34.49 | 23.9 | 994.48 | 35.014 | 31.881 | 996.231 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.04 | 0.07
994.53 | 34.77 | 23.99 | 34.5 | 23.92 | 994.45 | 35.019 | 31.894 | 996.199 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | 0.08 | 0.07
994.53 | 34.76 | 23.98 | 34.49 | 23.91 | 994.48 | 35.014 | 31.892 | 996.228 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | 0.05 | 0.07
994.52 | 34.78 | 23.98 | 34.51 | 23.96 | 994.47 | 35.024 | 31.936 | 996.214 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | 0.05 | 0.02
994.52 | 34.79 | 23.97 | 34.51 | 23.93 | 994.48 | 35.03 | 31.911 | 996.227 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.04 | 0.04
994.52 | 34.8 | 23.99 | 34.55 | 23.95 | 994.46 | 35.062 | 31.926 | 996.21 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.06 | 0.04
994.51 | 34.8 | 24.01 | 34.52 | 23.98 | 994.47 | 35.035 | 31.959 | 996.218 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.04 | 0.03
994.52 | 34.8 | 24 | 34.55 | 24 | 994.47 | 35.062 | 31.98 | 996.218 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.05 | 0
994.51 | 34.81 | 24.02 | 34.56 | 23.94 | 994.47 | 35.072 | 31.915 | 996.221 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.04 | 0.08
994.51 | 34.81 | 24.04 | 34.56 | 23.95 | 994.45 | 35.072 | 31.932 | 996.193 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.06 | 0.09
994.52 | 34.82 | 24.02 | 34.54 | 24.06 | 994.45 | 35.051 | 32.037 | 996.195 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.07 | -0.04
994.52 | 34.82 | 24.04 | 34.55 | 23.99 | 994.48 | 35.062 | 31.972 | 996.222 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | 0.04 | 0.05
994.51 | 34.83 | 24.05 | 34.56 | 24.01 | 994.46 | 35.072 | 31.985 | 996.205 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | 0.05 | 0.04
994.52 | 34.81 | 24.05 | 34.55 | 23.98 | 994.46 | 35.062 | 31.957 | 996.207 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0.06 | 0.07
994.52 | 34.84 | 24.1 | 34.56 | 23.99 | 994.46 | 35.072 | 31.966 | 996.209 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.06 | 0.11
994.5 | 34.83 | 24.06 | 34.58 | 24.04 | 994.47 | 35.083 | 32.014 | 996.218 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.03 | 0.02
994.52 | 34.83 | 24.07 | 34.56 | 24.05 | 994.44 | 35.072 | 32.027 | 996.19 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | 0.08 | 0.02
994.52 | 34.83 | 24.07 | 34.61 | 24.01 | 994.46 | 35.115 | 31.991 | 996.208 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.06 | 0.06
994.51 | 34.86 | 24.09 | 34.64 | 24.01 | 994.47 | 35.137 | 31.989 | 996.216 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.04 | 0.08
994.5 | 34.85 | 24.1 | 34.6 | 24.04 | 994.45 | 35.105 | 32.016 | 996.195 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.05 | 0.06
994.51 | 34.86 | 24.11 | 34.6 | 24.05 | 994.44 | 35.105 | 32.024 | 996.188 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0.07 | 0.06
994.5 | 34.88 | 24.11 | 34.61 | 24.04 | 994.44 | 35.115 | 32.014 | 996.191 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | 0.06 | 0.07
994.5 | 34.88 | 24.09 | 34.66 | 24.01 | 994.45 | 35.153 | 31.985 | 996.192 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.05 | 0.08
994.49 | 34.89 | 24.1 | 34.64 | 24.07 | 994.45 | 35.137 | 32.05 | 996.199 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.04 | 0.03
994.48 | 34.87 | 24.15 | 34.65 | 24.08 | 994.44 | 35.142 | 32.056 | 996.187 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.04 | 0.07
994.48 | 34.88 | 24.13 | 34.64 | 24.07 | 994.43 | 35.137 | 32.045 | 996.176 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.05 | 0.06
994.48 | 34.91 | 24.13 | 34.66 | 24.08 | 994.42 | 35.158 | 32.058 | 996.165 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.06 | 0.05
994.48 | 34.89 | 24.15 | 34.65 | 24.07 | 994.41 | 35.147 | 32.045 | 996.157 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.07 | 0.08
994.49 | 34.91 | 24.15 | 34.66 | 24.12 | 994.42 | 35.158 | 32.092 | 996.17 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.07 | 0.03
994.49 | 34.9 | 24.16 | 34.64 | 24.1 | 994.43 | 35.137 | 32.073 | 996.175 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0.06 | 0.06
994.5 | 34.91 | 24.15 | 34.67 | 24.12 | 994.42 | 35.163 | 32.092 | 996.171 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.08 | 0.03
994.5 | 34.91 | 24.18 | 34.68 | 24.08 | 994.45 | 35.169 | 32.058 | 996.192 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.05 | 0.1
994.51 | 34.9 | 24.18 | 34.66 | 24.15 | 994.45 | 35.153 | 32.123 | 996.196 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.06 | 0.03
994.52 | 34.92 | 24.19 | 34.68 | 24.12 | 994.45 | 35.174 | 32.102 | 996.196 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.07 | 0.07
994.54 | 34.93 | 24.2 | 34.67 | 24.13 | 994.46 | 35.163 | 32.109 | 996.209 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0.08 | 0.07
994.52 | 34.94 | 24.21 | 34.72 | 24.12 | 994.47 | 35.206 | 32.096 | 996.218 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.05 | 0.09
994.52 | 34.94 | 24.2 | 34.7 | 24.14 | 994.47 | 35.19 | 32.115 | 996.214 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.05 | 0.06
994.52 | 34.93 | 24.21 | 34.69 | 24.14 | 994.47 | 35.179 | 32.119 | 996.216 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.05 | 0.07
994.52 | 34.95 | 24.19 | 34.74 | 24.14 | 994.48 | 35.222 | 32.121 | 996.226 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.04 | 0.05
994.53 | 34.95 | 24.22 | 34.75 | 24.12 | 994.47 | 35.227 | 32.092 | 996.213 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.06 | 0.1
994.53 | 34.95 | 24.21 | 34.72 | 24.16 | 994.48 | 35.206 | 32.134 | 996.223 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.05 | 0.05
994.54 | 34.96 | 24.24 | 34.75 | 24.16 | 994.47 | 35.233 | 32.134 | 996.218 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.07 | 0.08
994.55 | 34.97 | 24.23 | 34.74 | 24.15 | 994.5 | 35.222 | 32.128 | 996.242 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.05 | 0.08
994.55 | 34.97 | 24.24 | 34.77 | 24.19 | 994.51 | 35.243 | 32.167 | 996.253 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.04 | 0.05
994.56 | 34.97 | 24.29 | 34.72 | 24.23 | 994.51 | 35.206 | 32.203 | 996.258 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.05 | 0.06
994.57 | 34.97 | 24.28 | 34.7 | 24.27 | 994.52 | 35.19 | 32.245 | 996.265 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | 0.05 | 0.01
994.56 | 34.97 | 24.28 | 34.75 | 24.21 | 994.54 | 35.233 | 32.186 | 996.288 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.02 | 0.07
994.56 | 34.97 | 24.31 | 34.75 | 24.28 | 994.52 | 35.233 | 32.26 | 996.27 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.04 | 0.03
994.55 | 34.99 | 24.3 | 34.73 | 24.28 | 994.52 | 35.217 | 32.258 | 996.265 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0.03 | 0.02
994.57 | 34.99 | 24.27 | 34.77 | 24.22 | 994.51 | 35.243 | 32.193 | 996.256 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.06 | 0.05
994.56 | 34.96 | 24.32 | 34.77 | 24.21 | 994.53 | 35.249 | 32.19 | 996.278 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.03 | 0.11
994.57 | 34.99 | 24.31 | 34.75 | 24.22 | 994.51 | 35.227 | 32.197 | 996.259 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.06 | 0.09
994.56 | 34.98 | 24.3 | 34.79 | 24.22 | 994.52 | 35.259 | 32.195 | 996.266 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.04 | 0.08
994.54 | 34.97 | 24.33 | 34.79 | 24.28 | 994.51 | 35.259 | 32.254 | 996.26 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.03 | 0.05
994.55 | 35 | 24.32 | 34.77 | 24.22 | 994.5 | 35.243 | 32.199 | 996.242 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.05 | 0.1
994.54 | 35.01 | 24.34 | 34.79 | 24.24 | 994.51 | 35.265 | 32.218 | 996.255 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.03 | 0.1
994.55 | 35 | 24.33 | 34.77 | 24.29 | 994.49 | 35.243 | 32.264 | 996.236 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.06 | 0.04
994.54 | 35.01 | 24.34 | 34.81 | 24.23 | 994.5 | 35.281 | 32.21 | 996.247 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.04 | 0.11
994.54 | 35 | 24.35 | 34.78 | 24.25 | 994.47 | 35.254 | 32.231 | 996.214 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.07 | 0.1
994.52 | 35.02 | 24.34 | 34.79 | 24.28 | 994.48 | 35.259 | 32.254 | 996.223 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.04 | 0.06
994.54 | 35.04 | 24.38 | 34.8 | 24.27 | 994.48 | 35.27 | 32.249 | 996.226 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.06 | 0.11
994.54 | 34.99 | 24.39 | 34.8 | 24.31 | 994.49 | 35.27 | 32.291 | 996.235 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.05 | 0.08
994.55 | 35.04 | 24.37 | 34.81 | 24.33 | 994.48 | 35.281 | 32.312 | 996.23 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.07 | 0.04
994.54 | 35.03 | 24.37 | 34.85 | 24.28 | 994.5 | 35.313 | 32.258 | 996.242 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.04 | 0.09
994.53 | 35.04 | 24.37 | 34.82 | 24.3 | 994.47 | 35.286 | 32.275 | 996.22 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.06 | 0.07
994.54 | 35.05 | 24.36 | 34.83 | 24.32 | 994.49 | 35.297 | 32.3 | 996.234 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.05 | 0.04
994.55 | 35.05 | 24.39 | 34.86 | 24.26 | 994.48 | 35.318 | 32.235 | 996.227 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.07 | 0.13
994.56 | 35.06 | 24.39 | 34.83 | 24.31 | 994.47 | 35.297 | 32.287 | 996.219 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.09 | 0.08
994.55 | 35.06 | 24.38 | 34.83 | 24.31 | 994.5 | 35.297 | 32.287 | 996.251 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.05 | 0.07
994.55 | 35.05 | 24.41 | 34.89 | 24.25 | 994.47 | 35.345 | 32.231 | 996.22 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.08 | 0.16
994.56 | 35.05 | 24.41 | 34.83 | 24.32 | 994.48 | 35.297 | 32.298 | 996.227 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.08 | 0.09
994.57 | 35.03 | 24.44 | 34.82 | 24.28 | 994.5 | 35.291 | 32.262 | 996.248 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.07 | 0.16
994.58 | 35.06 | 24.45 | 34.8 | 24.35 | 994.49 | 35.275 | 32.327 | 996.241 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0.09 | 0.1
994.58 | 35.04 | 24.48 | 34.84 | 24.36 | 994.51 | 35.307 | 32.342 | 996.26 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.07 | 0.12
994.58 | 35.05 | 24.44 | 34.81 | 24.4 | 994.52 | 35.281 | 32.378 | 996.269 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.06 | 0.04
994.59 | 35.06 | 24.46 | 34.87 | 24.38 | 994.51 | 35.329 | 32.359 | 996.26 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.08 | 0.08
994.59 | 35.06 | 24.44 | 34.85 | 24.41 | 994.53 | 35.313 | 32.384 | 996.278 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.06 | 0.03
994.6 | 35.08 | 24.46 | 34.88 | 24.31 | 994.52 | 35.34 | 32.283 | 996.271 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.08 | 0.15
994.61 | 35.08 | 24.44 | 34.89 | 24.38 | 994.53 | 35.35 | 32.361 | 996.275 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.08 | 0.06
994.61 | 35.1 | 24.46 | 34.9 | 24.33 | 994.54 | 35.356 | 32.31 | 996.282 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.07 | 0.13
994.62 | 35.1 | 24.49 | 34.89 | 24.34 | 994.55 | 35.35 | 32.315 | 996.293 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.07 | 0.15
994.63 | 35.08 | 24.5 | 34.89 | 24.39 | 994.55 | 35.345 | 32.369 | 996.296 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.08 | 0.11
994.63 | 35.09 | 24.48 | 34.87 | 24.38 | 994.57 | 35.334 | 32.361 | 996.32 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.06 | 0.1
994.66 | 35.09 | 24.49 | 34.89 | 24.4 | 994.58 | 35.345 | 32.38 | 996.323 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.08 | 0.09
994.67 | 35.1 | 24.5 | 34.88 | 24.42 | 994.59 | 35.34 | 32.397 | 996.341 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.08 | 0.08
994.68 | 35.11 | 24.49 | 34.89 | 24.41 | 994.62 | 35.35 | 32.384 | 996.365 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.06 | 0.08
994.69 | 35.11 | 24.51 | 34.9 | 24.38 | 994.63 | 35.356 | 32.357 | 996.375 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.06 | 0.13
994.71 | 35.1 | 24.52 | 34.89 | 24.42 | 994.62 | 35.35 | 32.394 | 996.369 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.09 | 0.1
994.72 | 35.12 | 24.51 | 34.92 | 24.47 | 994.65 | 35.372 | 32.445 | 996.398 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.07 | 0.04
994.73 | 35.12 | 24.53 | 34.89 | 24.47 | 994.65 | 35.345 | 32.445 | 996.392 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.08 | 0.06
994.75 | 35.11 | 24.57 | 34.92 | 24.43 | 994.68 | 35.372 | 32.407 | 996.423 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.07 | 0.14
994.76 | 35.09 | 24.56 | 34.89 | 24.44 | 994.71 | 35.35 | 32.42 | 996.453 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.05 | 0.12
994.77 | 35.1 | 24.56 | 34.91 | 24.48 | 994.72 | 35.361 | 32.455 | 996.464 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.05 | 0.08
994.78 | 35.11 | 24.55 | 34.94 | 24.45 | 994.71 | 35.388 | 32.426 | 996.456 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.07 | 0.1
994.78 | 35.11 | 24.59 | 34.92 | 24.45 | 994.72 | 35.372 | 32.43 | 996.469 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.06 | 0.14
994.77 | 35.12 | 24.58 | 34.92 | 24.44 | 994.73 | 35.377 | 32.413 | 996.477 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.04 | 0.14
994.77 | 35.12 | 24.58 | 34.93 | 24.45 | 994.73 | 35.382 | 32.43 | 996.472 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.04 | 0.13
994.77 | 35.14 | 24.59 | 34.92 | 24.48 | 994.72 | 35.372 | 32.455 | 996.464 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.05 | 0.11
994.78 | 35.13 | 24.58 | 34.95 | 24.45 | 994.71 | 35.398 | 32.426 | 996.458 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.07 | 0.13
994.79 | 35.15 | 24.6 | 34.96 | 24.48 | 994.72 | 35.404 | 32.457 | 996.463 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.07 | 0.12
994.79 | 35.14 | 24.59 | 34.94 | 24.48 | 994.73 | 35.388 | 32.46 | 996.476 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.06 | 0.11
994.8 | 35.14 | 24.58 | 34.95 | 24.45 | 994.72 | 35.398 | 32.428 | 996.47 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.08 | 0.13
994.81 | 35.17 | 24.6 | 35 | 24.46 | 994.74 | 35.441 | 32.441 | 996.487 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.07 | 0.14
994.81 | 35.14 | 24.62 | 34.97 | 24.53 | 994.74 | 35.414 | 32.504 | 996.488 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.07 | 0.09
994.79 | 35.16 | 24.61 | 34.97 | 24.51 | 994.74 | 35.414 | 32.487 | 996.49 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.05 | 0.1
994.78 | 35.17 | 24.58 | 34.99 | 24.45 | 994.73 | 35.436 | 32.422 | 996.479 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.05 | 0.13
994.79 | 35.16 | 24.62 | 35.01 | 24.48 | 994.72 | 35.452 | 32.458 | 996.463 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.14
994.78 | 35.17 | 24.63 | 34.98 | 24.54 | 994.73 | 35.42 | 32.521 | 996.482 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.05 | 0.09
994.78 | 35.18 | 24.67 | 34.99 | 24.53 | 994.72 | 35.436 | 32.504 | 996.463 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.06 | 0.14
994.78 | 35.16 | 24.63 | 34.97 | 24.57 | 994.72 | 35.414 | 32.546 | 996.468 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.06 | 0.06
994.79 | 35.18 | 24.62 | 34.97 | 24.48 | 994.73 | 35.414 | 32.458 | 996.472 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.06 | 0.14
994.8 | 35.18 | 24.63 | 35.03 | 24.48 | 994.74 | 35.468 | 32.455 | 996.482 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.15
994.82 | 35.17 | 24.66 | 34.98 | 24.6 | 994.76 | 35.425 | 32.577 | 996.505 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.06 | 0.06
994.82 | 35.18 | 24.65 | 35 | 24.5 | 994.77 | 35.441 | 32.481 | 996.516 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.05 | 0.15
994.82 | 35.18 | 24.65 | 34.99 | 24.56 | 994.77 | 35.43 | 32.535 | 996.516 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.05 | 0.09
994.82 | 35.19 | 24.66 | 34.99 | 24.55 | 994.77 | 35.436 | 32.527 | 996.519 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.05 | 0.11
994.82 | 35.19 | 24.63 | 35.03 | 24.54 | 994.76 | 35.462 | 32.521 | 996.507 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.09
994.83 | 35.18 | 24.66 | 35.03 | 24.48 | 994.76 | 35.468 | 32.455 | 996.502 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.18
994.84 | 35.19 | 24.65 | 35.01 | 24.54 | 994.78 | 35.452 | 32.519 | 996.529 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.06 | 0.11
994.83 | 35.19 | 24.64 | 35.04 | 24.5 | 994.79 | 35.473 | 32.474 | 996.539 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.04 | 0.14
994.83 | 35.22 | 24.67 | 35 | 24.53 | 994.78 | 35.441 | 32.512 | 996.524 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.05 | 0.14
994.83 | 35.21 | 24.67 | 35.05 | 24.51 | 994.77 | 35.478 | 32.485 | 996.519 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.16
994.85 | 35.21 | 24.7 | 35.03 | 24.55 | 994.78 | 35.468 | 32.525 | 996.529 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.07 | 0.15
994.85 | 35.2 | 24.73 | 35.02 | 24.64 | 994.8 | 35.457 | 32.615 | 996.542 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.05 | 0.09
994.85 | 35.21 | 24.71 | 35.03 | 24.61 | 994.8 | 35.468 | 32.59 | 996.547 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.05 | 0.1
994.86 | 35.19 | 24.74 | 35.03 | 24.54 | 994.8 | 35.468 | 32.521 | 996.545 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.2
994.86 | 35.2 | 24.74 | 35.03 | 24.63 | 994.8 | 35.468 | 32.603 | 996.547 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.06 | 0.11
994.85 | 35.21 | 24.72 | 35.05 | 24.58 | 994.79 | 35.478 | 32.561 | 996.54 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.14
994.85 | 35.19 | 24.75 | 34.99 | 24.64 | 994.8 | 35.436 | 32.617 | 996.549 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.05 | 0.11
994.85 | 35.21 | 24.74 | 35.03 | 24.63 | 994.79 | 35.468 | 32.607 | 996.538 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.06 | 0.11
994.86 | 35.21 | 24.74 | 35.03 | 24.58 | 994.76 | 35.462 | 32.556 | 996.511 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.1 | 0.16
994.89 | 35.21 | 24.76 | 35.01 | 24.58 | 994.81 | 35.452 | 32.561 | 996.553 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.08 | 0.18
994.88 | 35.23 | 24.75 | 35.05 | 24.55 | 994.83 | 35.478 | 32.531 | 996.572 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.05 | 0.2
994.87 | 35.22 | 24.75 | 35.08 | 24.57 | 994.81 | 35.505 | 32.544 | 996.557 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.18
994.87 | 35.22 | 24.77 | 35.06 | 24.61 | 994.82 | 35.489 | 32.586 | 996.567 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.05 | 0.16
994.89 | 35.23 | 24.77 | 35.07 | 24.62 | 994.81 | 35.5 | 32.6 | 996.561 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.08 | 0.15
994.89 | 35.23 | 24.76 | 35.06 | 24.59 | 994.82 | 35.489 | 32.569 | 996.571 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.07 | 0.17
994.89 | 35.21 | 24.79 | 35.05 | 24.59 | 994.82 | 35.484 | 32.563 | 996.571 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.07 | 0.2
994.9 | 35.23 | 24.79 | 35.05 | 24.63 | 994.85 | 35.478 | 32.609 | 996.598 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.05 | 0.16
994.91 | 35.22 | 24.81 | 35.08 | 24.64 | 994.84 | 35.505 | 32.619 | 996.589 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.17
994.92 | 35.22 | 24.82 | 35.03 | 24.67 | 994.83 | 35.462 | 32.643 | 996.581 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.09 | 0.15
994.93 | 35.24 | 24.8 | 35.06 | 24.69 | 994.85 | 35.489 | 32.672 | 996.6 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.08 | 0.11
994.92 | 35.24 | 24.8 | 35.07 | 24.67 | 994.86 | 35.5 | 32.651 | 996.609 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.06 | 0.13
994.92 | 35.24 | 24.81 | 35.08 | 24.65 | 994.86 | 35.505 | 32.624 | 996.605 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.16
994.92 | 35.24 | 24.81 | 35.11 | 24.68 | 994.84 | 35.532 | 32.661 | 996.588 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.13
994.92 | 35.25 | 24.83 | 35.1 | 24.66 | 994.84 | 35.521 | 32.638 | 996.587 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.08 | 0.17
994.92 | 35.24 | 24.85 | 35.06 | 24.73 | 994.85 | 35.489 | 32.705 | 996.598 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.07 | 0.12
994.92 | 35.24 | 24.86 | 35.1 | 24.68 | 994.85 | 35.521 | 32.661 | 996.6 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.18
994.91 | 35.24 | 24.86 | 35.06 | 24.7 | 994.85 | 35.489 | 32.676 | 996.595 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.06 | 0.16
994.9 | 35.25 | 24.83 | 35.09 | 24.66 | 994.84 | 35.516 | 32.642 | 996.591 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.17
994.9 | 35.25 | 24.84 | 35.06 | 24.71 | 994.86 | 35.489 | 32.687 | 996.607 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.04 | 0.13
994.91 | 35.23 | 24.89 | 35.07 | 24.68 | 994.84 | 35.5 | 32.653 | 996.59 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.07 | 0.21
994.91 | 35.26 | 24.87 | 35.05 | 24.79 | 994.86 | 35.484 | 32.769 | 996.603 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.05 | 0.08
994.91 | 35.27 | 24.86 | 35.1 | 24.72 | 994.85 | 35.521 | 32.697 | 996.593 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.06 | 0.14
994.92 | 35.27 | 24.9 | 35.12 | 24.7 | 994.86 | 35.542 | 32.68 | 996.604 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.2
994.94 | 35.26 | 24.9 | 35.09 | 24.72 | 994.87 | 35.516 | 32.697 | 996.618 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.07 | 0.18
994.95 | 35.24 | 24.89 | 35.08 | 24.76 | 994.89 | 35.51 | 32.737 | 996.633 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.13
994.95 | 35.26 | 24.9 | 35.09 | 24.77 | 994.9 | 35.516 | 32.743 | 996.648 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.05 | 0.13
994.94 | 35.25 | 24.89 | 35.1 | 24.73 | 994.9 | 35.521 | 32.712 | 996.65 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.04 | 0.16
994.96 | 35.25 | 24.89 | 35.14 | 24.73 | 994.89 | 35.558 | 32.705 | 996.639 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.07 | 0.16
994.97 | 35.26 | 24.89 | 35.11 | 24.7 | 994.91 | 35.532 | 32.674 | 996.656 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.19
994.96 | 35.25 | 24.91 | 35.11 | 24.74 | 994.91 | 35.532 | 32.716 | 996.658 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.05 | 0.17
994.96 | 35.29 | 24.9 | 35.12 | 24.74 | 994.88 | 35.542 | 32.716 | 996.63 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.08 | 0.16
994.97 | 35.27 | 24.91 | 35.12 | 24.76 | 994.9 | 35.542 | 32.733 | 996.65 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.15
994.98 | 35.27 | 24.95 | 35.11 | 24.76 | 994.93 | 35.532 | 32.737 | 996.678 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.05 | 0.19
994.98 | 35.27 | 24.94 | 35.12 | 24.8 | 994.93 | 35.542 | 32.773 | 996.673 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.05 | 0.14
994.99 | 35.27 | 24.93 | 35.12 | 24.77 | 994.93 | 35.537 | 32.746 | 996.678 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.16
995 | 35.29 | 24.92 | 35.12 | 24.76 | 994.95 | 35.542 | 32.733 | 996.697 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.05 | 0.16
995.02 | 35.27 | 24.93 | 35.14 | 24.74 | 994.94 | 35.558 | 32.722 | 996.683 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.19
995.03 | 35.28 | 24.94 | 35.12 | 24.76 | 994.98 | 35.542 | 32.737 | 996.723 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.05 | 0.18
995.04 | 35.28 | 24.96 | 35.15 | 24.73 | 994.98 | 35.569 | 32.705 | 996.729 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.23
995.04 | 35.28 | 24.96 | 35.12 | 24.8 | 994.99 | 35.542 | 32.781 | 996.733 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.05 | 0.16
995.04 | 35.3 | 24.94 | 35.13 | 24.78 | 994.99 | 35.553 | 32.758 | 996.741 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.05 | 0.16
995.05 | 35.29 | 24.96 | 35.15 | 24.77 | 994.99 | 35.569 | 32.746 | 996.736 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.19
995.05 | 35.3 | 24.97 | 35.13 | 24.8 | 995 | 35.548 | 32.781 | 996.748 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.05 | 0.17
995.05 | 35.29 | 24.97 | 35.14 | 24.8 | 995 | 35.558 | 32.777 | 996.747 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.05 | 0.17
995.05 | 35.28 | 24.98 | 35.11 | 24.79 | 994.99 | 35.532 | 32.766 | 996.733 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.06 | 0.19
995.05 | 35.3 | 24.97 | 35.15 | 24.79 | 994.99 | 35.564 | 32.764 | 996.737 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.18
995.04 | 35.29 | 24.99 | 35.13 | 24.82 | 994.99 | 35.553 | 32.8 | 996.738 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.05 | 0.17
995.02 | 35.26 | 25.01 | 35.13 | 24.82 | 994.98 | 35.553 | 32.798 | 996.729 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.04 | 0.19
995.02 | 35.29 | 25.01 | 35.14 | 24.8 | 994.96 | 35.558 | 32.777 | 996.708 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.21
995.02 | 35.3 | 25 | 35.14 | 24.86 | 994.96 | 35.558 | 32.836 | 996.704 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.14
995.02 | 35.29 | 25.04 | 35.15 | 24.84 | 994.95 | 35.569 | 32.817 | 996.695 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.2
995.02 | 35.29 | 25.03 | 35.15 | 24.83 | 994.95 | 35.569 | 32.808 | 996.692 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.2
995.03 | 35.28 | 25.04 | 35.17 | 24.82 | 994.95 | 35.58 | 32.802 | 996.692 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.08 | 0.22
995.03 | 35.31 | 25.01 | 35.16 | 24.85 | 994.97 | 35.575 | 32.83 | 996.719 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.16
995.03 | 35.31 | 25.01 | 35.15 | 24.81 | 994.96 | 35.569 | 32.79 | 996.708 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.07 | 0.2
995.03 | 35.31 | 25.02 | 35.14 | 24.81 | 994.95 | 35.558 | 32.788 | 996.699 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.08 | 0.21
995.02 | 35.31 | 25.06 | 35.19 | 24.84 | 994.94 | 35.596 | 32.815 | 996.688 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.08 | 0.22
995.02 | 35.28 | 25.06 | 35.17 | 24.85 | 994.96 | 35.58 | 32.828 | 996.706 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.06 | 0.21
995.01 | 35.32 | 25.05 | 35.14 | 24.9 | 994.96 | 35.558 | 32.874 | 996.705 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.05 | 0.15
995 | 35.32 | 25.05 | 35.17 | 24.79 | 994.93 | 35.58 | 32.766 | 996.674 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.26
995 | 35.31 | 25.07 | 35.15 | 24.86 | 994.95 | 35.564 | 32.836 | 996.699 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.05 | 0.21
994.98 | 35.32 | 25.07 | 35.15 | 24.84 | 994.91 | 35.569 | 32.819 | 996.66 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.07 | 0.23
994.96 | 35.31 | 25.09 | 35.17 | 24.86 | 994.9 | 35.585 | 32.838 | 996.647 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.23
994.93 | 35.31 | 25.08 | 35.15 | 24.88 | 994.88 | 35.569 | 32.853 | 996.624 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.05 | 0.2
994.91 | 35.3 | 25.1 | 35.16 | 24.9 | 994.85 | 35.575 | 32.878 | 996.592 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.2
994.92 | 35.32 | 25.09 | 35.17 | 24.86 | 994.83 | 35.58 | 32.832 | 996.579 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.09 | 0.23
994.92 | 35.32 | 25.09 | 35.2 | 24.86 | 994.84 | 35.607 | 32.84 | 996.582 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.08 | 0.23
994.91 | 35.31 | 25.1 | 35.16 | 24.92 | 994.84 | 35.575 | 32.897 | 996.582 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.18
994.91 | 35.32 | 25.11 | 35.17 | 24.91 | 994.84 | 35.58 | 32.884 | 996.59 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.2
994.9 | 35.34 | 25.11 | 35.17 | 24.9 | 994.82 | 35.58 | 32.872 | 996.571 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.08 | 0.21
994.89 | 35.33 | 25.13 | 35.19 | 24.87 | 994.82 | 35.601 | 32.851 | 996.565 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.26
994.88 | 35.3 | 25.14 | 35.15 | 24.92 | 994.82 | 35.569 | 32.901 | 996.565 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.22
994.88 | 35.31 | 25.14 | 35.19 | 24.93 | 994.81 | 35.596 | 32.911 | 996.552 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.07 | 0.21
994.87 | 35.32 | 25.15 | 35.17 | 24.98 | 994.8 | 35.58 | 32.954 | 996.549 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.17
994.87 | 35.32 | 25.16 | 35.17 | 24.95 | 994.8 | 35.58 | 32.928 | 996.543 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.21
994.87 | 35.32 | 25.15 | 35.15 | 24.94 | 994.79 | 35.569 | 32.912 | 996.532 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.08 | 0.21
994.88 | 35.33 | 25.16 | 35.2 | 24.94 | 994.79 | 35.607 | 32.916 | 996.538 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.09 | 0.22
994.88 | 35.32 | 25.19 | 35.2 | 24.96 | 994.8 | 35.607 | 32.933 | 996.545 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.08 | 0.23
994.88 | 35.34 | 25.16 | 35.19 | 24.97 | 994.8 | 35.596 | 32.945 | 996.545 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.08 | 0.19
994.88 | 35.33 | 25.17 | 35.19 | 24.95 | 994.81 | 35.596 | 32.926 | 996.557 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.22
994.87 | 35.34 | 25.17 | 35.17 | 24.95 | 994.8 | 35.58 | 32.926 | 996.545 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.07 | 0.22
994.87 | 35.34 | 25.18 | 35.2 | 24.97 | 994.78 | 35.607 | 32.945 | 996.531 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.09 | 0.21
994.87 | 35.34 | 25.2 | 35.19 | 24.9 | 994.79 | 35.596 | 32.878 | 996.541 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.08 | 0.3
994.86 | 35.34 | 25.18 | 35.2 | 24.95 | 994.79 | 35.607 | 32.93 | 996.532 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.23
994.86 | 35.33 | 25.19 | 35.22 | 24.95 | 994.77 | 35.623 | 32.928 | 996.518 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.09 | 0.24
994.87 | 35.33 | 25.2 | 35.2 | 24.99 | 994.78 | 35.612 | 32.972 | 996.53 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.09 | 0.21
994.87 | 35.34 | 25.21 | 35.19 | 25 | 994.78 | 35.596 | 32.979 | 996.525 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.09 | 0.21
994.88 | 35.34 | 25.21 | 35.2 | 24.99 | 994.78 | 35.607 | 32.97 | 996.529 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.1 | 0.22
994.88 | 35.33 | 25.21 | 35.19 | 25.03 | 994.8 | 35.596 | 33.004 | 996.548 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.08 | 0.18
994.88 | 35.34 | 25.24 | 35.17 | 24.97 | 994.78 | 35.585 | 32.949 | 996.529 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.1 | 0.27
994.88 | 35.35 | 25.22 | 35.2 | 25.03 | 994.79 | 35.612 | 33.012 | 996.539 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.09 | 0.19
994.89 | 35.36 | 25.24 | 35.2 | 25.01 | 994.8 | 35.607 | 32.989 | 996.547 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.09 | 0.23
994.9 | 35.35 | 25.24 | 35.2 | 25.01 | 994.8 | 35.612 | 32.989 | 996.546 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.1 | 0.23
994.91 | 35.34 | 25.24 | 35.2 | 25.03 | 994.81 | 35.612 | 33.012 | 996.56 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.1 | 0.21
994.91 | 35.32 | 25.25 | 35.19 | 25.06 | 994.83 | 35.596 | 33.04 | 996.573 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.19
994.93 | 35.35 | 25.25 | 35.21 | 25.05 | 994.83 | 35.617 | 33.029 | 996.577 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.1 | 0.2
994.95 | 35.35 | 25.26 | 35.21 | 24.99 | 994.85 | 35.617 | 32.972 | 996.597 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.1 | 0.27
994.96 | 35.35 | 25.27 | 35.22 | 25.06 | 994.86 | 35.623 | 33.042 | 996.609 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.1 | 0.21
994.97 | 35.35 | 25.28 | 35.22 | 25.05 | 994.87 | 35.623 | 33.023 | 996.621 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.1 | 0.23
994.97 | 35.34 | 25.29 | 35.2 | 25.05 | 994.87 | 35.607 | 33.027 | 996.621 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.1 | 0.24
994.96 | 35.32 | 25.3 | 35.21 | 25.05 | 994.88 | 35.617 | 33.029 | 996.624 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.08 | 0.25
994.95 | 35.34 | 25.3 | 35.18 | 25.06 | 994.88 | 35.591 | 33.033 | 996.631 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.07 | 0.24
994.95 | 35.36 | 25.29 | 35.2 | 25.08 | 994.86 | 35.612 | 33.057 | 996.61 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.09 | 0.21
994.94 | 35.35 | 25.31 | 35.21 | 25.06 | 994.87 | 35.617 | 33.038 | 996.614 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.25
994.93 | 35.34 | 25.32 | 35.18 | 25.08 | 994.86 | 35.591 | 33.055 | 996.605 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.07 | 0.24
994.93 | 35.34 | 25.32 | 35.2 | 25.05 | 994.86 | 35.607 | 33.027 | 996.602 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.27
994.93 | 35.36 | 25.32 | 35.22 | 25.1 | 994.85 | 35.628 | 33.076 | 996.597 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.08 | 0.22
994.93 | 35.36 | 25.33 | 35.22 | 25.08 | 994.84 | 35.623 | 33.054 | 996.589 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.09 | 0.25
994.93 | 35.35 | 25.34 | 35.22 | 25.1 | 994.85 | 35.623 | 33.082 | 996.599 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.24
994.94 | 35.35 | 25.35 | 35.18 | 25.15 | 994.85 | 35.591 | 33.126 | 996.6 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.09 | 0.2
994.94 | 35.35 | 25.34 | 35.24 | 25.08 | 994.86 | 35.639 | 33.059 | 996.604 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.08 | 0.26
994.94 | 35.36 | 25.34 | 35.21 | 25.13 | 994.86 | 35.617 | 33.107 | 996.604 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.08 | 0.21
994.94 | 35.36 | 25.35 | 35.2 | 25.09 | 994.85 | 35.607 | 33.071 | 996.601 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.09 | 0.26
994.95 | 35.36 | 25.36 | 35.22 | 25.08 | 994.86 | 35.623 | 33.061 | 996.61 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.09 | 0.28
994.96 | 35.34 | 25.37 | 35.2 | 25.13 | 994.86 | 35.607 | 33.107 | 996.609 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.1 | 0.24
994.98 | 35.35 | 25.35 | 35.22 | 25.12 | 994.89 | 35.623 | 33.101 | 996.632 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.09 | 0.23
995 | 35.35 | 25.38 | 35.24 | 25.14 | 994.89 | 35.644 | 33.118 | 996.633 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.11 | 0.24
995.02 | 35.37 | 25.38 | 35.22 | 25.14 | 994.92 | 35.628 | 33.12 | 996.662 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.1 | 0.24
995.04 | 35.35 | 25.39 | 35.22 | 25.14 | 994.94 | 35.623 | 33.12 | 996.685 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.1 | 0.25
995.05 | 35.35 | 25.4 | 35.23 | 25.15 | 994.97 | 35.633 | 33.13 | 996.719 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.08 | 0.25
995.05 | 35.37 | 25.4 | 35.23 | 25.15 | 994.98 | 35.633 | 33.13 | 996.723 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.25
995.05 | 35.36 | 25.4 | 35.23 | 25.17 | 994.97 | 35.633 | 33.151 | 996.718 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.23
995.04 | 35.36 | 25.4 | 35.23 | 25.16 | 994.98 | 35.633 | 33.137 | 996.727 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.24
995.04 | 35.36 | 25.41 | 35.2 | 25.16 | 994.98 | 35.607 | 33.136 | 996.728 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.25
995.04 | 35.36 | 25.43 | 35.24 | 25.19 | 994.98 | 35.644 | 33.162 | 996.727 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.24
995.05 | 35.37 | 25.42 | 35.21 | 25.17 | 994.97 | 35.617 | 33.149 | 996.718 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.08 | 0.25
995.06 | 35.38 | 25.41 | 35.23 | 25.17 | 994.98 | 35.633 | 33.143 | 996.73 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.08 | 0.24
995.06 | 35.39 | 25.43 | 35.26 | 25.15 | 994.98 | 35.655 | 33.126 | 996.725 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.28
995.07 | 35.38 | 25.42 | 35.22 | 25.16 | 994.99 | 35.628 | 33.137 | 996.736 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.08 | 0.26
995.08 | 35.36 | 25.43 | 35.24 | 25.19 | 995 | 35.639 | 33.168 | 996.75 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.08 | 0.24
995.09 | 35.38 | 25.43 | 35.23 | 25.15 | 995.02 | 35.633 | 33.13 | 996.767 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.28
995.09 | 35.39 | 25.42 | 35.23 | 25.2 | 995.04 | 35.633 | 33.179 | 996.785 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.05 | 0.22
995.09 | 35.38 | 25.43 | 35.25 | 25.19 | 995.02 | 35.649 | 33.162 | 996.763 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.24
995.1 | 35.38 | 25.45 | 35.25 | 25.22 | 995.02 | 35.649 | 33.2 | 996.765 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.23
995.11 | 35.36 | 25.46 | 35.27 | 25.21 | 995.02 | 35.665 | 33.191 | 996.769 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.09 | 0.25
995.11 | 35.38 | 25.47 | 35.2 | 25.22 | 995.04 | 35.607 | 33.193 | 996.783 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.07 | 0.25
995.12 | 35.37 | 25.48 | 35.23 | 25.21 | 995.05 | 35.633 | 33.189 | 996.793 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.27
995.13 | 35.38 | 25.5 | 35.24 | 25.21 | 995.06 | 35.644 | 33.189 | 996.802 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.29
995.14 | 35.37 | 25.49 | 35.23 | 25.23 | 995.08 | 35.633 | 33.21 | 996.827 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.26
995.15 | 35.38 | 25.49 | 35.24 | 25.24 | 995.09 | 35.644 | 33.221 | 996.832 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.25
995.16 | 35.36 | 25.51 | 35.23 | 25.24 | 995.11 | 35.633 | 33.214 | 996.853 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.05 | 0.27
995.16 | 35.37 | 25.51 | 35.25 | 25.26 | 995.11 | 35.649 | 33.233 | 996.852 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.05 | 0.25
995.15 | 35.37 | 25.52 | 35.22 | 25.23 | 995.12 | 35.628 | 33.208 | 996.864 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.03 | 0.29
995.15 | 35.37 | 25.51 | 35.26 | 25.23 | 995.1 | 35.655 | 33.204 | 996.843 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.05 | 0.28
995.15 | 35.38 | 25.51 | 35.26 | 25.26 | 995.08 | 35.66 | 33.24 | 996.827 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.07 | 0.25
995.15 | 35.4 | 25.5 | 35.27 | 25.21 | 995.07 | 35.671 | 33.191 | 996.818 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.29
995.13 | 35.39 | 25.5 | 35.27 | 25.23 | 995.07 | 35.671 | 33.21 | 996.816 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.27
995.13 | 35.41 | 25.51 | 35.26 | 25.23 | 995.07 | 35.66 | 33.212 | 996.815 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.28
995.12 | 35.39 | 25.52 | 35.25 | 25.24 | 995.06 | 35.649 | 33.216 | 996.804 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.28
995.1 | 35.4 | 25.53 | 35.23 | 25.26 | 995.04 | 35.633 | 33.237 | 996.789 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.06 | 0.27
995.1 | 35.39 | 25.53 | 35.26 | 25.27 | 995.04 | 35.66 | 33.246 | 996.782 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.26
995.1 | 35.38 | 25.53 | 35.24 | 25.27 | 995.02 | 35.644 | 33.252 | 996.769 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.08 | 0.26
995.09 | 35.38 | 25.55 | 35.24 | 25.3 | 995.02 | 35.644 | 33.279 | 996.764 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.25
995.09 | 35.39 | 25.55 | 35.24 | 25.27 | 995.02 | 35.644 | 33.244 | 996.765 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.28
995.09 | 35.4 | 25.55 | 35.25 | 25.27 | 995.01 | 35.649 | 33.244 | 996.755 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.08 | 0.28
995.09 | 35.39 | 25.56 | 35.26 | 25.28 | 995.01 | 35.66 | 33.258 | 996.758 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.28
995.09 | 35.38 | 25.56 | 35.26 | 25.31 | 995.03 | 35.655 | 33.29 | 996.774 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.25
995.08 | 35.4 | 25.58 | 35.27 | 25.28 | 995.02 | 35.665 | 33.261 | 996.769 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.3
995.08 | 35.4 | 25.57 | 35.25 | 25.3 | 995 | 35.649 | 33.279 | 996.751 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.08 | 0.27
995.08 | 35.42 | 25.58 | 35.24 | 25.3 | 995 | 35.644 | 33.279 | 996.751 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.08 | 0.28
995.09 | 35.38 | 25.57 | 35.25 | 25.31 | 995.02 | 35.649 | 33.29 | 996.767 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.26
995.08 | 35.38 | 25.58 | 35.28 | 25.3 | 995.02 | 35.676 | 33.273 | 996.764 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.06 | 0.28
995.07 | 35.39 | 25.59 | 35.25 | 25.31 | 995.01 | 35.649 | 33.286 | 996.756 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.28
995.06 | 35.38 | 25.61 | 35.22 | 25.32 | 995 | 35.628 | 33.298 | 996.745 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.29
995.06 | 35.37 | 25.61 | 35.26 | 25.33 | 995 | 35.655 | 33.309 | 996.744 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.06 | 0.28
995.07 | 35.4 | 25.61 | 35.23 | 25.32 | 995 | 35.633 | 33.298 | 996.751 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.07 | 0.29
995.08 | 35.38 | 25.62 | 35.26 | 25.33 | 994.99 | 35.66 | 33.305 | 996.74 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.09 | 0.29
995.09 | 35.38 | 25.64 | 35.28 | 25.34 | 995.01 | 35.676 | 33.317 | 996.758 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.08 | 0.3
995.1 | 35.38 | 25.64 | 35.27 | 25.36 | 995.03 | 35.665 | 33.336 | 996.773 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.07 | 0.28
995.1 | 35.38 | 25.64 | 35.26 | 25.37 | 995.03 | 35.66 | 33.345 | 996.773 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.07 | 0.27
995.1 | 35.39 | 25.63 | 35.25 | 25.33 | 995.03 | 35.649 | 33.303 | 996.779 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.3
995.09 | 35.39 | 25.66 | 35.28 | 25.35 | 995.03 | 35.676 | 33.324 | 996.778 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.06 | 0.31
995.07 | 35.39 | 25.66 | 35.26 | 25.38 | 995.03 | 35.655 | 33.353 | 996.779 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.04 | 0.28
995.05 | 35.37 | 25.64 | 35.25 | 25.39 | 995.02 | 35.649 | 33.364 | 996.763 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.03 | 0.25
995.04 | 35.38 | 25.64 | 35.26 | 25.34 | 994.98 | 35.66 | 33.321 | 996.725 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.3
995.03 | 35.4 | 25.65 | 35.26 | 25.38 | 994.96 | 35.655 | 33.353 | 996.71 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.27
995.06 | 35.39 | 25.67 | 35.25 | 25.37 | 994.97 | 35.649 | 33.345 | 996.72 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.09 | 0.3
995.09 | 35.4 | 25.66 | 35.26 | 25.38 | 994.99 | 35.655 | 33.353 | 996.735 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.1 | 0.28
995.14 | 35.41 | 25.66 | 35.24 | 25.36 | 995.02 | 35.639 | 33.34 | 996.764 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.12 | 0.3
995.16 | 35.41 | 25.66 | 35.26 | 25.37 | 995.06 | 35.66 | 33.351 | 996.805 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.1 | 0.29
995.15 | 35.39 | 25.68 | 35.28 | 25.37 | 995.09 | 35.676 | 33.351 | 996.835 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.06 | 0.31
995.1 | 35.39 | 25.69 | 35.27 | 25.41 | 995.11 | 35.671 | 33.384 | 996.853 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | -0.01 | 0.28
995.04 | 35.39 | 25.7 | 35.23 | 25.44 | 995.03 | 35.633 | 33.416 | 996.777 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.01 | 0.26
995.02 | 35.4 | 25.7 | 35.25 | 25.45 | 994.97 | 35.649 | 33.427 | 996.715 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.05 | 0.25
995 | 35.37 | 25.71 | 35.23 | 25.41 | 994.94 | 35.633 | 33.385 | 996.69 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.3
994.99 | 35.38 | 25.73 | 35.25 | 25.42 | 994.94 | 35.649 | 33.393 | 996.691 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.05 | 0.31
995.01 | 35.37 | 25.74 | 35.24 | 25.45 | 994.94 | 35.644 | 33.429 | 996.682 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.29
995.02 | 35.37 | 25.75 | 35.26 | 25.43 | 994.92 | 35.655 | 33.405 | 996.671 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.1 | 0.32
995.04 | 35.37 | 25.74 | 35.23 | 25.44 | 994.96 | 35.633 | 33.416 | 996.704 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.08 | 0.3
995.06 | 35.39 | 25.73 | 35.24 | 25.43 | 994.98 | 35.639 | 33.41 | 996.723 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.08 | 0.3
995.08 | 35.38 | 25.75 | 35.23 | 25.4 | 994.98 | 35.633 | 33.378 | 996.73 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.1 | 0.35
995.1 | 35.37 | 25.75 | 35.26 | 25.46 | 995.01 | 35.66 | 33.435 | 996.758 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.09 | 0.29
995.12 | 35.38 | 25.74 | 35.25 | 25.47 | 995.03 | 35.649 | 33.448 | 996.777 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.09 | 0.27
995.13 | 35.38 | 25.77 | 35.25 | 25.46 | 995.06 | 35.649 | 33.439 | 996.802 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.31
995.13 | 35.37 | 25.76 | 35.26 | 25.47 | 995.06 | 35.655 | 33.446 | 996.809 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.07 | 0.29
995.13 | 35.4 | 25.76 | 35.23 | 25.43 | 995.06 | 35.633 | 33.41 | 996.811 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.07 | 0.33
995.12 | 35.39 | 25.75 | 35.26 | 25.46 | 995.06 | 35.655 | 33.439 | 996.804 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.29
995.1 | 35.42 | 25.75 | 35.27 | 25.45 | 995.06 | 35.671 | 33.431 | 996.805 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.04 | 0.3
995.09 | 35.39 | 25.76 | 35.26 | 25.45 | 995.03 | 35.66 | 33.427 | 996.772 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.31
995.09 | 35.38 | 25.77 | 35.25 | 25.45 | 995.03 | 35.649 | 33.424 | 996.775 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.32
995.11 | 35.37 | 25.77 | 35.26 | 25.44 | 995.02 | 35.66 | 33.42 | 996.771 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.09 | 0.33
995.11 | 35.39 | 25.77 | 35.24 | 25.48 | 995.03 | 35.644 | 33.456 | 996.78 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.08 | 0.29
995.12 | 35.38 | 25.77 | 35.28 | 25.48 | 995.04 | 35.676 | 33.462 | 996.783 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.08 | 0.29
995.12 | 35.4 | 25.77 | 35.27 | 25.48 | 995.06 | 35.665 | 33.46 | 996.81 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.29
995.11 | 35.42 | 25.77 | 35.27 | 25.47 | 995.06 | 35.665 | 33.443 | 996.802 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.05 | 0.3
995.11 | 35.41 | 25.79 | 35.27 | 25.47 | 995.07 | 35.671 | 33.446 | 996.819 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.04 | 0.32
995.12 | 35.39 | 25.8 | 35.26 | 25.5 | 995.06 | 35.655 | 33.477 | 996.802 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.3
995.14 | 35.39 | 25.8 | 35.26 | 25.51 | 995.08 | 35.655 | 33.488 | 996.823 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.29
995.15 | 35.4 | 25.79 | 35.27 | 25.48 | 995.09 | 35.671 | 33.458 | 996.832 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.31
995.16 | 35.41 | 25.78 | 35.26 | 25.43 | 995.1 | 35.66 | 33.41 | 996.851 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.35
995.18 | 35.41 | 25.8 | 35.27 | 25.51 | 995.1 | 35.665 | 33.49 | 996.851 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.08 | 0.29
995.17 | 35.41 | 25.8 | 35.26 | 25.55 | 995.13 | 35.655 | 33.528 | 996.874 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.04 | 0.25
995.15 | 35.41 | 25.81 | 35.26 | 25.55 | 995.1 | 35.655 | 33.528 | 996.844 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.05 | 0.26
995.13 | 35.39 | 25.82 | 35.25 | 25.48 | 995.08 | 35.649 | 33.456 | 996.828 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.05 | 0.34
995.11 | 35.41 | 25.83 | 35.27 | 25.49 | 995.06 | 35.665 | 33.466 | 996.811 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.05 | 0.34
995.11 | 35.39 | 25.84 | 35.24 | 25.49 | 995.04 | 35.644 | 33.467 | 996.79 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.35
995.11 | 35.4 | 25.84 | 35.26 | 25.5 | 995.04 | 35.66 | 33.479 | 996.786 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.34
995.12 | 35.4 | 25.83 | 35.29 | 25.5 | 995.05 | 35.681 | 33.477 | 996.794 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.07 | 0.33
995.12 | 35.4 | 25.83 | 35.24 | 25.52 | 995.06 | 35.639 | 33.492 | 996.806 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.31
995.12 | 35.4 | 25.84 | 35.27 | 25.49 | 995.07 | 35.665 | 33.466 | 996.813 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.05 | 0.35
995.12 | 35.39 | 25.85 | 35.25 | 25.52 | 995.06 | 35.649 | 33.5 | 996.804 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.33
995.12 | 35.38 | 25.85 | 35.26 | 25.57 | 995.06 | 35.655 | 33.548 | 996.804 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.28
995.12 | 35.38 | 25.86 | 35.28 | 25.52 | 995.07 | 35.676 | 33.502 | 996.821 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.05 | 0.34
995.12 | 35.4 | 25.86 | 35.22 | 25.55 | 995.05 | 35.623 | 33.525 | 996.793 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.07 | 0.31
995.11 | 35.4 | 25.86 | 35.27 | 25.54 | 995.04 | 35.665 | 33.513 | 996.785 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.32
995.1 | 35.39 | 25.86 | 35.26 | 25.59 | 995.04 | 35.655 | 33.567 | 996.787 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.27
995.1 | 35.42 | 25.85 | 35.28 | 25.53 | 995.04 | 35.676 | 33.509 | 996.786 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.32
995.08 | 35.4 | 25.85 | 35.27 | 25.5 | 995.02 | 35.671 | 33.477 | 996.771 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.35
995.07 | 35.41 | 25.85 | 35.27 | 25.52 | 995 | 35.665 | 33.494 | 996.75 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.33
995.05 | 35.42 | 25.84 | 35.27 | 25.54 | 995 | 35.671 | 33.521 | 996.744 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.05 | 0.3
995.04 | 35.41 | 25.86 | 35.29 | 25.54 | 994.98 | 35.681 | 33.521 | 996.722 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.32
995.03 | 35.41 | 25.86 | 35.25 | 25.59 | 994.96 | 35.649 | 33.567 | 996.707 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.07 | 0.27
995.02 | 35.41 | 25.86 | 35.27 | 25.55 | 994.97 | 35.665 | 33.53 | 996.712 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.05 | 0.31
995.02 | 35.42 | 25.87 | 35.29 | 25.56 | 994.97 | 35.687 | 33.538 | 996.716 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.05 | 0.31
995.03 | 35.42 | 25.87 | 35.26 | 25.6 | 994.96 | 35.66 | 33.576 | 996.705 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.07 | 0.27
995.04 | 35.41 | 25.89 | 35.27 | 25.56 | 994.98 | 35.665 | 33.538 | 996.73 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.33
995.06 | 35.43 | 25.88 | 35.26 | 25.59 | 994.99 | 35.66 | 33.565 | 996.735 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.07 | 0.29
995.06 | 35.42 | 25.88 | 35.28 | 25.57 | 995 | 35.676 | 33.544 | 996.75 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.31
995.06 | 35.42 | 25.86 | 35.3 | 25.53 | 994.99 | 35.692 | 33.508 | 996.74 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.07 | 0.33
995.06 | 35.43 | 25.88 | 35.28 | 25.57 | 995 | 35.676 | 33.548 | 996.746 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.31
995.03 | 35.42 | 25.89 | 35.28 | 25.53 | 995 | 35.676 | 33.509 | 996.745 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.03 | 0.36
995.02 | 35.41 | 25.91 | 35.28 | 25.59 | 994.98 | 35.676 | 33.569 | 996.729 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.04 | 0.32
995.01 | 35.41 | 25.91 | 35.24 | 25.62 | 994.98 | 35.644 | 33.595 | 996.724 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.03 | 0.29
995 | 35.39 | 25.91 | 35.27 | 25.58 | 994.96 | 35.665 | 33.559 | 996.708 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.04 | 0.33
995 | 35.41 | 25.91 | 35.28 | 25.55 | 994.96 | 35.676 | 33.53 | 996.704 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.04 | 0.36
995 | 35.43 | 25.9 | 35.26 | 25.58 | 994.95 | 35.655 | 33.557 | 996.696 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.05 | 0.32
994.99 | 35.41 | 25.93 | 35.26 | 25.61 | 994.95 | 35.66 | 33.584 | 996.698 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.04 | 0.32
995 | 35.42 | 25.9 | 35.25 | 25.57 | 994.95 | 35.649 | 33.546 | 996.694 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.05 | 0.33
995 | 35.42 | 25.9 | 35.28 | 25.55 | 994.95 | 35.676 | 33.525 | 996.698 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.05 | 0.35
994.99 | 35.43 | 25.91 | 35.28 | 25.58 | 994.93 | 35.676 | 33.561 | 996.681 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.33
994.97 | 35.42 | 25.93 | 35.3 | 25.58 | 994.92 | 35.692 | 33.561 | 996.668 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.05 | 0.35
994.95 | 35.4 | 25.93 | 35.27 | 25.6 | 994.91 | 35.665 | 33.572 | 996.652 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.04 | 0.33
994.94 | 35.41 | 25.93 | 35.28 | 25.6 | 994.9 | 35.676 | 33.572 | 996.643 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.04 | 0.33
994.93 | 35.41 | 25.93 | 35.27 | 25.63 | 994.88 | 35.665 | 33.603 | 996.627 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.05 | 0.3
994.91 | 35.41 | 25.93 | 35.28 | 25.59 | 994.87 | 35.676 | 33.567 | 996.614 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.04 | 0.34
994.9 | 35.42 | 25.93 | 35.28 | 25.56 | 994.85 | 35.676 | 33.538 | 996.6 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.05 | 0.37
994.89 | 35.4 | 25.95 | 35.3 | 25.58 | 994.84 | 35.692 | 33.561 | 996.588 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.05 | 0.37
994.87 | 35.4 | 25.95 | 35.27 | 25.66 | 994.82 | 35.671 | 33.637 | 996.568 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.05 | 0.29
994.88 | 35.41 | 25.95 | 35.27 | 25.6 | 994.81 | 35.665 | 33.58 | 996.558 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.35
994.89 | 35.42 | 25.95 | 35.26 | 25.59 | 994.83 | 35.655 | 33.565 | 996.574 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.36
994.91 | 35.42 | 25.95 | 35.29 | 25.57 | 994.83 | 35.687 | 33.549 | 996.581 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.38
994.93 | 35.41 | 25.96 | 35.26 | 25.61 | 994.86 | 35.66 | 33.59 | 996.609 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.35
994.94 | 35.42 | 25.95 | 35.28 | 25.59 | 994.87 | 35.676 | 33.569 | 996.614 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.36
994.95 | 35.41 | 25.96 | 35.27 | 25.61 | 994.88 | 35.671 | 33.584 | 996.622 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.35
994.96 | 35.41 | 25.97 | 35.27 | 25.62 | 994.89 | 35.665 | 33.595 | 996.636 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.35
994.97 | 35.42 | 25.97 | 35.27 | 25.57 | 994.9 | 35.671 | 33.549 | 996.65 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.4
994.98 | 35.43 | 25.96 | 35.26 | 25.64 | 994.9 | 35.66 | 33.612 | 996.649 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.08 | 0.32
994.99 | 35.43 | 25.95 | 35.29 | 25.64 | 994.92 | 35.687 | 33.62 | 996.671 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.31
995 | 35.44 | 25.96 | 35.28 | 25.6 | 994.93 | 35.676 | 33.578 | 996.677 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.07 | 0.36
994.99 | 35.44 | 25.95 | 35.3 | 25.62 | 994.93 | 35.692 | 33.597 | 996.68 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.33
994.99 | 35.45 | 25.95 | 35.31 | 25.57 | 994.92 | 35.703 | 33.546 | 996.669 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.38
994.99 | 35.44 | 25.93 | 35.28 | 25.57 | 994.93 | 35.676 | 33.549 | 996.676 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.36
994.98 | 35.44 | 25.93 | 35.29 | 25.57 | 994.93 | 35.687 | 33.544 | 996.677 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.05 | 0.36
994.97 | 35.45 | 25.94 | 35.32 | 25.6 | 994.91 | 35.708 | 33.576 | 996.654 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.34
994.97 | 35.45 | 25.96 | 35.31 | 25.61 | 994.9 | 35.703 | 33.588 | 996.645 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.35
994.99 | 35.46 | 25.96 | 35.31 | 25.59 | 994.92 | 35.697 | 33.57 | 996.665 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.37
995.01 | 35.46 | 25.97 | 35.33 | 25.62 | 994.94 | 35.719 | 33.593 | 996.689 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.35
995.03 | 35.43 | 25.99 | 35.27 | 25.61 | 994.95 | 35.671 | 33.591 | 996.701 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.08 | 0.38
995.05 | 35.44 | 25.97 | 35.33 | 25.6 | 994.99 | 35.713 | 33.58 | 996.734 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.06 | 0.37
995.07 | 35.46 | 25.96 | 35.32 | 25.59 | 994.99 | 35.708 | 33.565 | 996.734 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.08 | 0.37
995.08 | 35.46 | 25.97 | 35.33 | 25.6 | 995.01 | 35.713 | 33.572 | 996.755 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.37
995.09 | 35.45 | 25.97 | 35.31 | 25.64 | 995.03 | 35.703 | 33.612 | 996.779 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.33
995.1 | 35.47 | 25.96 | 35.32 | 25.63 | 995.03 | 35.708 | 33.605 | 996.777 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.33
995.12 | 35.46 | 25.97 | 35.33 | 25.61 | 995.05 | 35.713 | 33.588 | 996.794 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.36
995.13 | 35.47 | 25.98 | 35.33 | 25.64 | 995.06 | 35.719 | 33.612 | 996.803 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.34
995.14 | 35.46 | 25.98 | 35.31 | 25.65 | 995.06 | 35.697 | 33.63 | 996.805 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.08 | 0.33
995.15 | 35.45 | 25.98 | 35.31 | 25.6 | 995.06 | 35.697 | 33.582 | 996.804 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.09 | 0.38
995.16 | 35.46 | 25.99 | 35.33 | 25.59 | 995.09 | 35.719 | 33.563 | 996.835 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.4
995.18 | 35.47 | 25.98 | 35.33 | 25.62 | 995.09 | 35.713 | 33.599 | 996.839 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.09 | 0.36
995.19 | 35.48 | 25.99 | 35.33 | 25.61 | 995.11 | 35.713 | 33.588 | 996.861 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.08 | 0.38
995.2 | 35.46 | 26 | 35.32 | 25.62 | 995.13 | 35.708 | 33.599 | 996.875 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.38
995.21 | 35.47 | 25.98 | 35.32 | 25.61 | 995.13 | 35.708 | 33.591 | 996.878 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.08 | 0.37
995.2 | 35.48 | 25.98 | 35.33 | 25.57 | 995.14 | 35.713 | 33.551 | 996.887 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.41
995.21 | 35.46 | 25.99 | 35.36 | 25.58 | 995.13 | 35.74 | 33.561 | 996.879 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.08 | 0.41
995.19 | 35.47 | 25.98 | 35.34 | 25.6 | 995.13 | 35.724 | 33.572 | 996.88 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.38
995.19 | 35.47 | 25.99 | 35.32 | 25.58 | 995.12 | 35.708 | 33.557 | 996.87 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.41
995.17 | 35.46 | 26 | 35.34 | 25.6 | 995.12 | 35.724 | 33.576 | 996.865 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.05 | 0.4
995.16 | 35.47 | 25.99 | 35.33 | 25.63 | 995.1 | 35.713 | 33.611 | 996.846 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.36
995.15 | 35.47 | 25.99 | 35.34 | 25.6 | 995.09 | 35.724 | 33.582 | 996.833 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.39
995.15 | 35.47 | 26 | 35.35 | 25.62 | 995.09 | 35.735 | 33.593 | 996.838 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.38
995.16 | 35.47 | 26 | 35.34 | 25.62 | 995.09 | 35.729 | 33.601 | 996.835 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.38
995.16 | 35.48 | 26.01 | 35.31 | 25.62 | 995.1 | 35.703 | 33.593 | 996.845 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.06 | 0.39
995.15 | 35.46 | 26 | 35.32 | 25.65 | 995.1 | 35.708 | 33.624 | 996.848 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.05 | 0.35
995.15 | 35.47 | 26.01 | 35.33 | 25.6 | 995.1 | 35.719 | 33.576 | 996.851 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.05 | 0.41
995.15 | 35.46 | 26.01 | 35.36 | 25.6 | 995.09 | 35.74 | 33.572 | 996.839 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.06 | 0.41
995.15 | 35.47 | 26.01 | 35.33 | 25.64 | 995.1 | 35.713 | 33.622 | 996.843 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.05 | 0.37
995.15 | 35.47 | 26.01 | 35.36 | 25.63 | 995.09 | 35.74 | 33.603 | 996.832 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.06 | 0.38
995.15 | 35.46 | 26 | 35.36 | 25.63 | 995.08 | 35.74 | 33.609 | 996.83 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.07 | 0.37
995.14 | 35.49 | 26.02 | 35.35 | 25.66 | 995.09 | 35.735 | 33.637 | 996.833 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.05 | 0.36
995.13 | 35.46 | 26.02 | 35.34 | 25.59 | 995.08 | 35.724 | 33.57 | 996.827 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.05 | 0.43
995.12 | 35.47 | 26.02 | 35.37 | 25.61 | 995.06 | 35.751 | 33.586 | 996.809 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.06 | 0.41
995.12 | 35.46 | 26.03 | 35.34 | 25.63 | 995.07 | 35.724 | 33.603 | 996.813 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.05 | 0.4
995.11 | 35.47 | 26.03 | 35.33 | 25.69 | 995.07 | 35.713 | 33.666 | 996.814 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.04 | 0.34
995.12 | 35.47 | 26.03 | 35.33 | 25.64 | 995.06 | 35.713 | 33.616 | 996.809 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.39
995.14 | 35.47 | 26.02 | 35.33 | 25.62 | 995.07 | 35.719 | 33.595 | 996.82 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.4
995.14 | 35.48 | 26.04 | 35.38 | 25.64 | 995.09 | 35.756 | 33.62 | 996.833 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.05 | 0.4
995.16 | 35.47 | 26.04 | 35.38 | 25.63 | 995.08 | 35.756 | 33.609 | 996.828 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.08 | 0.41
995.17 | 35.46 | 26.04 | 35.33 | 25.66 | 995.09 | 35.713 | 33.639 | 996.841 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.38
995.17 | 35.49 | 26.03 | 35.34 | 25.61 | 995.1 | 35.724 | 33.588 | 996.849 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.07 | 0.42
995.19 | 35.47 | 26.03 | 35.35 | 25.63 | 995.11 | 35.735 | 33.607 | 996.861 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.08 | 0.4
995.2 | 35.47 | 26.04 | 35.33 | 25.64 | 995.13 | 35.719 | 33.618 | 996.873 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.07 | 0.4
995.18 | 35.47 | 26.04 | 35.32 | 25.66 | 995.14 | 35.708 | 33.635 | 996.886 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.04 | 0.38
995.17 | 35.47 | 26.04 | 35.34 | 25.63 | 995.12 | 35.729 | 33.607 | 996.865 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.05 | 0.41
995.16 | 35.48 | 26.03 | 35.34 | 25.64 | 995.12 | 35.729 | 33.614 | 996.869 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.04 | 0.39
995.14 | 35.48 | 26.04 | 35.37 | 25.64 | 995.1 | 35.751 | 33.62 | 996.843 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.04 | 0.4
995.13 | 35.49 | 26.02 | 35.38 | 25.66 | 995.07 | 35.761 | 33.639 | 996.818 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.06 | 0.36
995.13 | 35.47 | 26.06 | 35.34 | 25.63 | 995.06 | 35.724 | 33.607 | 996.805 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.43
995.13 | 35.47 | 26.06 | 35.29 | 25.67 | 995.07 | 35.687 | 33.649 | 996.815 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.06 | 0.39
995.15 | 35.48 | 26.05 | 35.34 | 25.64 | 995.07 | 35.724 | 33.618 | 996.816 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.08 | 0.41
995.16 | 35.47 | 26.04 | 35.35 | 25.66 | 995.09 | 35.735 | 33.633 | 996.835 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.07 | 0.38
995.17 | 35.49 | 26.04 | 35.36 | 25.64 | 995.12 | 35.74 | 33.62 | 996.862 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.05 | 0.4
995.17 | 35.49 | 26.04 | 35.33 | 25.64 | 995.11 | 35.713 | 33.618 | 996.855 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.06 | 0.4
995.17 | 35.49 | 26.04 | 35.36 | 25.59 | 995.12 | 35.74 | 33.567 | 996.868 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.05 | 0.45
995.16 | 35.47 | 26.03 | 35.36 | 25.66 | 995.12 | 35.74 | 33.633 | 996.868 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.04 | 0.37
995.15 | 35.48 | 26.05 | 35.33 | 25.65 | 995.1 | 35.719 | 33.63 | 996.843 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.05 | 0.4
995.13 | 35.47 | 26.05 | 35.36 | 25.66 | 995.09 | 35.74 | 33.633 | 996.836 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.04 | 0.39
995.13 | 35.47 | 26.07 | 35.35 | 25.62 | 995.08 | 35.735 | 33.593 | 996.825 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.05 | 0.45
995.13 | 35.47 | 26.05 | 35.36 | 25.68 | 995.08 | 35.74 | 33.656 | 996.822 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.05 | 0.37
995.13 | 35.48 | 26.05 | 35.36 | 25.65 | 995.07 | 35.74 | 33.624 | 996.821 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.4
995.14 | 35.48 | 26.05 | 35.33 | 25.63 | 995.08 | 35.719 | 33.611 | 996.822 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.42
995.15 | 35.48 | 26.06 | 35.36 | 25.61 | 995.08 | 35.745 | 33.586 | 996.831 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.07 | 0.45
995.16 | 35.47 | 26.05 | 35.38 | 25.62 | 995.09 | 35.761 | 33.597 | 996.833 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.07 | 0.43
995.17 | 35.47 | 26.05 | 35.36 | 25.65 | 995.09 | 35.74 | 33.624 | 996.832 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.08 | 0.4
995.18 | 35.48 | 26.05 | 35.36 | 25.65 | 995.12 | 35.745 | 33.631 | 996.868 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.4
995.2 | 35.48 | 26.05 | 35.36 | 25.63 | 995.12 | 35.74 | 33.603 | 996.867 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.08 | 0.42
995.2 | 35.51 | 26.04 | 35.36 | 25.68 | 995.14 | 35.74 | 33.654 | 996.885 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.06 | 0.36
995.2 | 35.51 | 26.03 | 35.37 | 25.6 | 995.14 | 35.751 | 33.58 | 996.883 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 0.43
995.2 | 35.51 | 26.03 | 35.38 | 25.6 | 995.13 | 35.761 | 33.58 | 996.875 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.43
995.2 | 35.51 | 26.04 | 35.39 | 25.6 | 995.14 | 35.767 | 33.576 | 996.891 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.44
995.2 | 35.51 | 26.04 | 35.38 | 25.66 | 995.13 | 35.761 | 33.639 | 996.872 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.38
995.21 | 35.5 | 26.03 | 35.39 | 25.59 | 995.14 | 35.772 | 33.563 | 996.884 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.07 | 0.44
995.23 | 35.51 | 26.04 | 35.38 | 25.62 | 995.15 | 35.761 | 33.595 | 996.894 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.42
995.23 | 35.5 | 26.03 | 35.39 | 25.61 | 995.18 | 35.772 | 33.588 | 996.924 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.05 | 0.42
995.23 | 35.51 | 26.05 | 35.39 | 25.63 | 995.17 | 35.772 | 33.607 | 996.917 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.42
995.24 | 35.51 | 26.05 | 35.38 | 25.64 | 995.18 | 35.761 | 33.616 | 996.924 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.41
995.23 | 35.48 | 26.05 | 35.38 | 25.61 | 995.18 | 35.761 | 33.591 | 996.93 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.05 | 0.44
995.23 | 35.49 | 26.07 | 35.38 | 25.66 | 995.16 | 35.761 | 33.635 | 996.91 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.07 | 0.41
995.22 | 35.5 | 26.06 | 35.39 | 25.61 | 995.17 | 35.767 | 33.586 | 996.912 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.05 | 0.45
995.23 | 35.51 | 26.03 | 35.39 | 25.64 | 995.16 | 35.767 | 33.622 | 996.909 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.07 | 0.39
995.23 | 35.51 | 26.04 | 35.39 | 25.6 | 995.16 | 35.772 | 33.574 | 996.91 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.07 | 0.44
995.22 | 35.51 | 26.04 | 35.38 | 25.63 | 995.17 | 35.761 | 33.603 | 996.92 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.05 | 0.41
995.22 | 35.51 | 26.03 | 35.39 | 25.61 | 995.16 | 35.767 | 33.591 | 996.911 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.42
995.22 | 35.51 | 26.05 | 35.42 | 25.61 | 995.17 | 35.793 | 33.59 | 996.912 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.05 | 0.44
995.23 | 35.53 | 26.03 | 35.4 | 25.64 | 995.17 | 35.777 | 33.616 | 996.912 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.39
995.22 | 35.52 | 26.04 | 35.41 | 25.63 | 995.16 | 35.783 | 33.605 | 996.909 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.06 | 0.41
995.21 | 35.52 | 26.03 | 35.4 | 25.62 | 995.15 | 35.777 | 33.599 | 996.898 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 0.41
995.22 | 35.51 | 26.04 | 35.39 | 25.59 | 995.15 | 35.767 | 33.567 | 996.893 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.07 | 0.45
995.22 | 35.53 | 26.03 | 35.4 | 25.6 | 995.16 | 35.777 | 33.574 | 996.905 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.43
995.24 | 35.55 | 26.04 | 35.42 | 25.62 | 995.16 | 35.793 | 33.599 | 996.908 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.08 | 0.42
995.25 | 35.54 | 26.02 | 35.43 | 25.63 | 995.18 | 35.799 | 33.607 | 996.926 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.07 | 0.39
995.27 | 35.53 | 26.02 | 35.43 | 25.62 | 995.19 | 35.804 | 33.601 | 996.937 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.08 | 0.4
995.28 | 35.54 | 26.02 | 35.43 | 25.57 | 995.2 | 35.804 | 33.548 | 996.944 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.08 | 0.45
995.29 | 35.54 | 26.03 | 35.41 | 25.62 | 995.22 | 35.783 | 33.593 | 996.968 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.41
995.3 | 35.55 | 26.02 | 35.44 | 25.59 | 995.22 | 35.809 | 33.565 | 996.963 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.08 | 0.43
995.3 | 35.56 | 26.03 | 35.45 | 25.6 | 995.23 | 35.82 | 33.572 | 996.976 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.07 | 0.43
995.31 | 35.56 | 26.01 | 35.45 | 25.59 | 995.22 | 35.815 | 33.567 | 996.967 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.09 | 0.42
995.31 | 35.56 | 26.01 | 35.46 | 25.58 | 995.24 | 35.826 | 33.555 | 996.99 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.07 | 0.43
995.3 | 35.56 | 26.01 | 35.48 | 25.6 | 995.25 | 35.842 | 33.58 | 996.995 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.05 | 0.41
995.3 | 35.56 | 26 | 35.48 | 25.57 | 995.23 | 35.842 | 33.551 | 996.977 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.07 | 0.43
995.31 | 35.56 | 26 | 35.49 | 25.6 | 995.24 | 35.852 | 33.576 | 996.983 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.07 | 0.07 | 0.4
995.31 | 35.57 | 26 | 35.46 | 25.57 | 995.23 | 35.826 | 33.549 | 996.975 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.08 | 0.43
995.33 | 35.58 | 25.99 | 35.48 | 25.59 | 995.24 | 35.842 | 33.565 | 996.99 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.09 | 0.4
995.33 | 35.59 | 25.99 | 35.46 | 25.58 | 995.24 | 35.831 | 33.561 | 996.991 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.09 | 0.41
995.33 | 35.58 | 25.98 | 35.46 | 25.54 | 995.25 | 35.831 | 33.515 | 996.993 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.08 | 0.44
995.32 | 35.6 | 25.99 | 35.48 | 25.58 | 995.25 | 35.842 | 33.561 | 996.993 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.07 | 0.41
995.31 | 35.58 | 25.98 | 35.49 | 25.55 | 995.24 | 35.852 | 33.528 | 996.988 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.07 | 0.43
995.32 | 35.59 | 25.98 | 35.48 | 25.56 | 995.23 | 35.847 | 33.534 | 996.977 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.09 | 0.42
995.33 | 35.6 | 25.95 | 35.49 | 25.57 | 995.25 | 35.852 | 33.544 | 996.998 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.08 | 0.38
995.34 | 35.61 | 25.94 | 35.53 | 25.52 | 995.25 | 35.884 | 33.502 | 996.996 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.09 | 0.42
995.33 | 35.64 | 25.91 | 35.51 | 25.51 | 995.27 | 35.868 | 33.487 | 997.015 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.06 | 0.4
995.35 | 35.67 | 25.87 | 35.5 | 25.5 | 995.26 | 35.858 | 33.473 | 997.005 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.09 | 0.37
995.36 | 35.73 | 25.82 | 35.48 | 25.52 | 995.27 | 35.847 | 33.502 | 997.021 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.09 | 0.3
995.37 | 35.75 | 25.78 | 35.53 | 25.5 | 995.3 | 35.884 | 33.479 | 997.042 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0.07 | 0.28
995.38 | 35.77 | 25.77 | 35.53 | 25.5 | 995.32 | 35.89 | 33.479 | 997.062 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.06 | 0.27
995.39 | 35.76 | 25.78 | 35.53 | 25.49 | 995.31 | 35.89 | 33.471 | 997.059 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.08 | 0.29
995.4 | 35.76 | 25.8 | 35.57 | 25.49 | 995.33 | 35.916 | 33.467 | 997.075 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.07 | 0.31
995.42 | 35.71 | 25.84 | 35.58 | 25.43 | 995.35 | 35.927 | 33.403 | 997.101 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.07 | 0.41
995.43 | 35.7 | 25.86 | 35.59 | 25.49 | 995.37 | 35.932 | 33.471 | 997.113 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.06 | 0.37
995.44 | 35.68 | 25.85 | 35.59 | 25.42 | 995.38 | 35.932 | 33.397 | 997.123 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.06 | 0.43
995.45 | 35.69 | 25.84 | 35.59 | 25.45 | 995.4 | 35.932 | 33.431 | 997.143 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.05 | 0.39
995.43 | 35.72 | 25.82 | 35.61 | 25.44 | 995.38 | 35.954 | 33.418 | 997.129 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.05 | 0.38
995.43 | 35.73 | 25.81 | 35.63 | 25.42 | 995.36 | 35.97 | 33.397 | 997.108 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.07 | 0.39
995.45 | 35.73 | 25.81 | 35.64 | 25.42 | 995.36 | 35.98 | 33.397 | 997.107 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.09 | 0.39
995.47 | 35.73 | 25.8 | 35.67 | 25.4 | 995.38 | 36.002 | 33.378 | 997.128 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0.09 | 0.4
995.47 | 35.74 | 25.79 | 35.64 | 25.4 | 995.39 | 35.98 | 33.38 | 997.138 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.08 | 0.39
995.46 | 35.74 | 25.78 | 35.66 | 25.37 | 995.4 | 35.996 | 33.351 | 997.148 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.06 | 0.41
995.45 | 35.76 | 25.77 | 35.66 | 25.37 | 995.41 | 35.991 | 33.343 | 997.153 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.04 | 0.4
995.45 | 35.79 | 25.75 | 35.69 | 25.32 | 995.4 | 36.023 | 33.3 | 997.145 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.05 | 0.43
995.45 | 35.8 | 25.73 | 35.7 | 25.3 | 995.39 | 36.028 | 33.275 | 997.138 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.06 | 0.43
995.43 | 35.81 | 25.72 | 35.73 | 25.28 | 995.39 | 36.055 | 33.256 | 997.134 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.04 | 0.44
995.44 | 35.81 | 25.7 | 35.71 | 25.31 | 995.38 | 36.034 | 33.282 | 997.126 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.06 | 0.39
995.44 | 35.83 | 25.69 | 35.74 | 25.27 | 995.37 | 36.061 | 33.242 | 997.121 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.07 | 0.42
995.44 | 35.83 | 25.68 | 35.77 | 25.3 | 995.36 | 36.087 | 33.273 | 997.108 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0.08 | 0.38
995.45 | 35.84 | 25.67 | 35.77 | 25.27 | 995.38 | 36.087 | 33.244 | 997.129 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.07 | 0.07 | 0.4
995.46 | 35.84 | 25.66 | 35.75 | 25.24 | 995.39 | 36.071 | 33.22 | 997.135 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.07 | 0.42
995.45 | 35.86 | 25.65 | 35.8 | 25.24 | 995.39 | 36.114 | 33.22 | 997.136 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0.06 | 0.41
995.47 | 35.89 | 25.63 | 35.81 | 25.23 | 995.39 | 36.125 | 33.206 | 997.134 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.08 | 0.4
995.47 | 35.9 | 25.62 | 35.83 | 25.17 | 995.4 | 36.141 | 33.151 | 997.147 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.07 | 0.07 | 0.45
995.48 | 35.89 | 25.61 | 35.83 | 25.18 | 995.4 | 36.141 | 33.155 | 997.15 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0.08 | 0.43
995.49 | 35.89 | 25.61 | 35.86 | 25.2 | 995.41 | 36.167 | 33.174 | 997.156 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.03 | 0.08 | 0.41
995.5 | 35.92 | 25.59 | 35.83 | 25.16 | 995.43 | 36.135 | 33.141 | 997.174 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.07 | 0.43
995.51 | 35.94 | 25.56 | 35.86 | 25.16 | 995.44 | 36.162 | 33.137 | 997.187 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.07 | 0.4
995.54 | 35.95 | 25.55 | 35.88 | 25.12 | 995.44 | 36.178 | 33.097 | 997.183 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.07 | 0.1 | 0.43
995.53 | 35.98 | 25.53 | 35.88 | 25.1 | 995.46 | 36.178 | 33.073 | 997.21 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.07 | 0.43
995.52 | 35.97 | 25.52 | 35.92 | 25.1 | 995.46 | 36.215 | 33.073 | 997.206 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0.06 | 0.42
995.53 | 36 | 25.49 | 35.94 | 25.06 | 995.46 | 36.231 | 33.033 | 997.202 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0.07 | 0.43
995.54 | 36 | 25.48 | 35.93 | 25.1 | 995.47 | 36.226 | 33.073 | 997.213 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.07 | 0.07 | 0.38
995.54 | 36 | 25.48 | 35.95 | 25.06 | 995.48 | 36.242 | 33.038 | 997.228 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0.06 | 0.42
995.52 | 36.01 | 25.47 | 35.96 | 25.03 | 995.48 | 36.247 | 33.012 | 997.223 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0.04 | 0.44
995.53 | 36.03 | 25.45 | 35.97 | 25.05 | 995.47 | 36.258 | 33.027 | 997.213 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0.06 | 0.4
995.54 | 36.04 | 25.42 | 36 | 25.02 | 995.48 | 36.285 | 32.993 | 997.229 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | 0.06 | 0.4
995.55 | 36.06 | 25.42 | 36.02 | 25.02 | 995.48 | 36.295 | 32.996 | 997.23 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | 0.07 | 0.4
995.56 | 36.06 | 25.4 | 36.04 | 25.02 | 995.51 | 36.312 | 33.002 | 997.258 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.02 | 0.05 | 0.38
995.57 | 36.1 | 25.38 | 36.04 | 24.97 | 995.51 | 36.312 | 32.943 | 997.261 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0.06 | 0.41
995.58 | 36.1 | 25.36 | 36.07 | 24.96 | 995.51 | 36.344 | 32.933 | 997.26 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.03 | 0.07 | 0.4
995.6 | 36.12 | 25.35 | 36.09 | 24.95 | 995.52 | 36.354 | 32.93 | 997.268 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.03 | 0.08 | 0.4
995.61 | 36.14 | 25.33 | 36.09 | 24.94 | 995.54 | 36.354 | 32.916 | 997.29 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0.07 | 0.39
995.61 | 36.14 | 25.31 | 36.09 | 24.92 | 995.56 | 36.36 | 32.901 | 997.307 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0.05 | 0.39
995.61 | 36.17 | 25.31 | 36.12 | 24.92 | 995.55 | 36.381 | 32.895 | 997.3 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0.06 | 0.39
995.61 | 36.16 | 25.3 | 36.15 | 24.91 | 995.56 | 36.408 | 32.886 | 997.304 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.01 | 0.05 | 0.39
995.63 | 36.18 | 25.3 | 36.16 | 24.86 | 995.55 | 36.413 | 32.84 | 997.299 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.02 | 0.08 | 0.44
995.64 | 36.18 | 25.28 | 36.19 | 24.85 | 995.57 | 36.44 | 32.83 | 997.32 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.01 | 0.07 | 0.43
995.65 | 36.21 | 25.26 | 36.19 | 24.9 | 995.58 | 36.44 | 32.876 | 997.322 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.02 | 0.07 | 0.36
995.66 | 36.19 | 25.24 | 36.2 | 24.82 | 995.61 | 36.445 | 32.8 | 997.359 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.01 | 0.05 | 0.42
995.7 | 36.22 | 25.23 | 36.19 | 24.82 | 995.62 | 36.44 | 32.8 | 997.366 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.03 | 0.08 | 0.41
995.7 | 36.27 | 25.2 | 36.27 | 24.83 | 995.65 | 36.504 | 32.809 | 997.399 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0 | 0.05 | 0.37
995.71 | 36.26 | 25.19 | 36.27 | 24.79 | 995.67 | 36.509 | 32.766 | 997.416 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.01 | 0.04 | 0.4
995.72 | 36.28 | 25.17 | 36.31 | 24.75 | 995.69 | 36.541 | 32.731 | 997.433 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.03 | 0.03 | 0.42
995.72 | 36.29 | 25.17 | 36.26 | 24.75 | 995.68 | 36.498 | 32.731 | 997.422 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.03 | 0.04 | 0.42
995.71 | 36.3 | 25.15 | 36.3 | 24.74 | 995.67 | 36.53 | 32.718 | 997.412 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0 | 0.04 | 0.41
995.71 | 36.32 | 25.13 | 36.32 | 24.74 | 995.64 | 36.552 | 32.718 | 997.391 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0 | 0.07 | 0.39
995.69 | 36.34 | 25.11 | 36.35 | 24.65 | 995.66 | 36.579 | 32.632 | 997.404 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.01 | 0.03 | 0.46
995.69 | 36.33 | 25.1 | 36.37 | 24.69 | 995.64 | 36.595 | 32.664 | 997.382 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.04 | 0.05 | 0.41
995.7 | 36.34 | 25.1 | 36.35 | 24.67 | 995.62 | 36.579 | 32.643 | 997.368 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.01 | 0.08 | 0.43
995.71 | 36.37 | 25.07 | 36.39 | 24.61 | 995.65 | 36.611 | 32.592 | 997.398 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.02 | 0.06 | 0.46
995.71 | 36.39 | 25.04 | 36.43 | 24.63 | 995.65 | 36.643 | 32.607 | 997.394 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.04 | 0.06 | 0.41
995.71 | 36.41 | 25.02 | 36.43 | 24.62 | 995.63 | 36.643 | 32.596 | 997.374 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.02 | 0.08 | 0.4
995.72 | 36.41 | 25.02 | 36.45 | 24.59 | 995.64 | 36.659 | 32.569 | 997.388 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.04 | 0.08 | 0.43
995.73 | 36.43 | 25 | 36.45 | 24.6 | 995.65 | 36.659 | 32.573 | 997.394 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.02 | 0.08 | 0.4
995.74 | 36.44 | 24.97 | 36.47 | 24.56 | 995.67 | 36.675 | 32.533 | 997.418 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.03 | 0.07 | 0.41
995.75 | 36.48 | 24.95 | 36.51 | 24.55 | 995.67 | 36.712 | 32.529 | 997.421 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.03 | 0.08 | 0.4
995.75 | 36.48 | 24.94 | 36.5 | 24.53 | 995.68 | 36.701 | 32.512 | 997.43 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.02 | 0.07 | 0.41
995.75 | 36.48 | 24.92 | 36.54 | 24.49 | 995.68 | 36.733 | 32.472 | 997.426 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.06 | 0.07 | 0.43
995.76 | 36.5 | 24.9 | 36.56 | 24.5 | 995.68 | 36.755 | 32.474 | 997.423 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.06 | 0.08 | 0.4
995.76 | 36.51 | 24.89 | 36.57 | 24.48 | 995.68 | 36.76 | 32.455 | 997.426 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.06 | 0.08 | 0.41
995.75 | 36.53 | 24.87 | 36.58 | 24.47 | 995.69 | 36.765 | 32.451 | 997.434 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.05 | 0.06 | 0.4
995.75 | 36.54 | 24.86 | 36.58 | 24.44 | 995.68 | 36.765 | 32.416 | 997.423 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.04 | 0.07 | 0.42
995.75 | 36.56 | 24.84 | 36.61 | 24.44 | 995.69 | 36.792 | 32.418 | 997.437 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.05 | 0.06 | 0.4
995.79 | 36.27 | 23.68 | 36.64 | 24.39 | 995.69 | 36.819 | 32.371 | 997.432 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.37 | 0.1 | -0.71
995.75 | 34.63 | 26.12 | 34.03 | 25.05 | 995.75 | 34.624 | 33.025 | 997.5 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.6 | 0 | 1.07
995.76 | 33.63 | 26.7 | 34.26 | 26.32 | 995.75 | 34.816 | 34.293 | 997.498 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.63 | 0.01 | 0.38
995.75 | 33.03 | 27.8 | 33.51 | 26.73 | 995.74 | 34.191 | 34.711 | 997.483 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.48 | 0.01 | 1.07
995.75 | 32.63 | 28.45 | 33.43 | 26.64 | 995.71 | 34.122 | 34.616 | 997.459 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.8 | 0.04 | 1.81
995.77 | 32.33 | 28.89 | 33.18 | 26.85 | 995.73 | 33.908 | 34.829 | 997.474 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.85 | 0.04 | 2.04
995.79 | 32.07 | 29.24 | 32.73 | 27.32 | 995.73 | 33.529 | 35.293 | 997.475 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.66 | 0.06 | 1.92
995.78 | 31.94 | 29.41 | 32.76 | 27.22 | 995.72 | 33.556 | 35.199 | 997.471 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.82 | 0.06 | 2.19
995.78 | 31.89 | 29.48 | 32.7 | 27.27 | 995.72 | 33.502 | 35.247 | 997.464 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.81 | 0.06 | 2.21
995.78 | 31.59 | 29.77 | 32.57 | 27.28 | 995.72 | 33.396 | 35.253 | 997.47 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.98 | 0.06 | 2.49
995.78 | 31.34 | 29.92 | 32.19 | 27.4 | 995.73 | 33.075 | 35.373 | 997.48 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.85 | 0.05 | 2.52
995.79 | 31.13 | 30.23 | 31.91 | 27.65 | 995.74 | 32.84 | 35.63 | 997.486 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.78 | 0.05 | 2.58
995.8 | 31.04 | 30.39 | 31.73 | 27.84 | 995.75 | 32.691 | 35.817 | 997.496 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.69 | 0.05 | 2.55
995.78 | 31 | 30.3 | 31.77 | 27.71 | 995.76 | 32.723 | 35.686 | 997.502 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.77 | 0.02 | 2.59
995.8 | 30.9 | 30.23 | 31.52 | 27.77 | 995.74 | 32.514 | 35.743 | 997.491 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.62 | 0.06 | 2.46
995.79 | 30.76 | 30.33 | 31.51 | 27.73 | 995.76 | 32.509 | 35.707 | 997.505 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.75 | 0.03 | 2.6
995.8 | 30.65 | 30.38 | 31.27 | 27.85 | 995.76 | 32.301 | 35.825 | 997.504 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.62 | 0.04 | 2.53
995.78 | 30.58 | 30.38 | 31.1 | 27.93 | 995.74 | 32.162 | 35.905 | 997.486 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.52 | 0.04 | 2.45
995.81 | 30.48 | 30.34 | 31.09 | 27.71 | 995.73 | 32.151 | 35.684 | 997.481 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.61 | 0.08 | 2.63
995.82 | 30.35 | 30.46 | 31.01 | 27.83 | 995.77 | 32.087 | 35.806 | 997.515 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.66 | 0.05 | 2.63
995.81 | 30.19 | 30.51 | 30.85 | 27.91 | 995.78 | 31.954 | 35.882 | 997.528 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.66 | 0.03 | 2.6
995.82 | 30.08 | 30.69 | 30.56 | 28 | 995.77 | 31.703 | 35.978 | 997.521 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.48 | 0.05 | 2.69
995.83 | 30.15 | 30.51 | 30.56 | 28.01 | 995.79 | 31.703 | 35.987 | 997.538 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.41 | 0.04 | 2.5
995.84 | 30.05 | 30.55 | 30.35 | 28 | 995.78 | 31.526 | 35.981 | 997.53 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.3 | 0.06 | 2.55
995.85 | 29.91 | 30.59 | 30.29 | 27.95 | 995.78 | 31.484 | 35.926 | 997.53 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.38 | 0.07 | 2.64
995.85 | 29.8 | 30.67 | 30.07 | 28.17 | 995.83 | 31.291 | 36.144 | 997.574 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.27 | 0.02 | 2.5
995.87 | 29.79 | 30.68 | 29.95 | 28.11 | 995.83 | 31.195 | 36.09 | 997.578 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.16 | 0.04 | 2.57
995.86 | 29.68 | 30.61 | 30.07 | 28.09 | 995.83 | 31.297 | 36.067 | 997.577 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.39 | 0.03 | 2.52
995.85 | 29.76 | 30.55 | 29.87 | 28.31 | 995.83 | 31.126 | 36.288 | 997.574 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.11 | 0.02 | 2.24
995.85 | 29.7 | 30.48 | 29.82 | 27.99 | 995.83 | 31.088 | 35.964 | 997.578 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.12 | 0.02 | 2.49
995.87 | 29.69 | 30.49 | 29.81 | 28.04 | 995.84 | 31.072 | 36.014 | 997.588 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.12 | 0.03 | 2.45
995.84 | 29.73 | 30.59 | 29.82 | 28.01 | 995.83 | 31.083 | 35.991 | 997.574 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.09 | 0.01 | 2.58
995.83 | 29.63 | 30.74 | 29.74 | 28.19 | 995.82 | 31.019 | 36.17 | 997.571 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.11 | 0.01 | 2.55
995.82 | 29.66 | 30.63 | 29.79 | 28.15 | 995.82 | 31.062 | 36.13 | 997.566 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.13 | 0 | 2.48
995.82 | 29.64 | 30.58 | 29.85 | 28 | 995.8 | 31.11 | 35.981 | 997.545 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.21 | 0.02 | 2.58
995.79 | 29.59 | 30.69 | 29.62 | 28.15 | 995.8 | 30.912 | 36.128 | 997.545 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.03 | -0.01 | 2.54
995.81 | 29.64 | 30.6 | 29.7 | 28.13 | 995.78 | 30.982 | 36.109 | 997.525 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.06 | 0.03 | 2.47
995.81 | 29.56 | 30.55 | 29.66 | 28.1 | 995.79 | 30.949 | 36.073 | 997.535 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.1 | 0.02 | 2.45
995.82 | 29.54 | 30.49 | 29.54 | 28.13 | 995.78 | 30.848 | 36.111 | 997.526 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0 | 0.04 | 2.36
995.82 | 29.51 | 30.5 | 29.4 | 28.29 | 995.82 | 30.731 | 36.269 | 997.563 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0 | 2.21
995.83 | 29.41 | 30.54 | 29.54 | 28.05 | 995.79 | 30.848 | 36.023 | 997.532 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.13 | 0.04 | 2.49
995.85 | 29.39 | 30.58 | 29.34 | 28.16 | 995.83 | 30.677 | 36.136 | 997.572 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0.02 | 2.42
995.83 | 29.54 | 30.95 | 29.43 | 28.14 | 995.82 | 30.757 | 36.113 | 997.567 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.01 | 2.81
995.83 | 29.81 | 30.82 | 29.77 | 27.99 | 995.81 | 31.04 | 35.964 | 997.561 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | 0.02 | 2.83
995.8 | 29.71 | 30.72 | 29.88 | 27.92 | 995.81 | 31.131 | 35.901 | 997.552 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.17 | -0.01 | 2.8
995.79 | 29.65 | 30.74 | 29.68 | 28.21 | 995.8 | 30.966 | 36.187 | 997.542 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.03 | -0.01 | 2.53
995.8 | 29.54 | 30.8 | 29.51 | 28.5 | 995.75 | 30.821 | 36.475 | 997.501 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.03 | 0.05 | 2.3
995.78 | 29.44 | 30.93 | 29.51 | 28.41 | 995.77 | 30.821 | 36.388 | 997.513 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.07 | 0.01 | 2.52
995.78 | 29.46 | 30.64 | 29.47 | 28.45 | 995.77 | 30.789 | 36.426 | 997.518 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.01 | 0.01 | 2.19
995.82 | 29.35 | 30.79 | 29.44 | 28.27 | 995.78 | 30.768 | 36.25 | 997.527 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.09 | 0.04 | 2.52
995.82 | 29.35 | 30.85 | 29.23 | 28.45 | 995.8 | 30.592 | 36.428 | 997.55 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.02 | 2.4
995.8 | 29.42 | 30.7 | 29.43 | 28.17 | 995.8 | 30.757 | 36.149 | 997.551 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.01 | 0 | 2.53
995.8 | 29.47 | 30.58 | 29.65 | 27.94 | 995.79 | 30.939 | 35.92 | 997.535 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.18 | 0.01 | 2.64
995.79 | 29.41 | 30.6 | 29.53 | 28 | 995.79 | 30.837 | 35.976 | 997.533 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.12 | 0 | 2.6
995.79 | 29.3 | 30.67 | 29.42 | 28.08 | 995.78 | 30.747 | 36.061 | 997.526 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.12 | 0.01 | 2.59
995.8 | 29.24 | 30.78 | 29.27 | 28.17 | 995.8 | 30.618 | 36.149 | 997.546 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.03 | 0 | 2.61
995.79 | 29.25 | 30.81 | 29.27 | 28.22 | 995.81 | 30.618 | 36.193 | 997.561 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.02 | -0.02 | 2.59
995.77 | 29.26 | 30.72 | 29.25 | 28.22 | 995.78 | 30.602 | 36.199 | 997.53 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.01 | -0.01 | 2.5
995.75 | 29.23 | 30.73 | 29.26 | 28.16 | 995.76 | 30.613 | 36.134 | 997.506 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.03 | -0.01 | 2.57
995.74 | 29.21 | 30.74 | 29.15 | 28.15 | 995.74 | 30.517 | 36.132 | 997.484 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0 | 2.59
995.74 | 29.19 | 30.65 | 29.21 | 28.17 | 995.72 | 30.57 | 36.144 | 997.464 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.02 | 0.02 | 2.48
995.73 | 29.15 | 30.72 | 29 | 28.37 | 995.74 | 30.394 | 36.35 | 997.489 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | -0.01 | 2.35
995.72 | 29.25 | 30.6 | 29.06 | 28.25 | 995.72 | 30.447 | 36.227 | 997.467 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0 | 2.35
995.7 | 29.18 | 30.6 | 29.09 | 28.13 | 995.71 | 30.474 | 36.107 | 997.455 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | -0.01 | 2.47
995.69 | 29.13 | 30.68 | 29.07 | 28.11 | 995.69 | 30.453 | 36.086 | 997.438 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0 | 2.57
995.68 | 29.14 | 30.62 | 29.09 | 28.08 | 995.68 | 30.474 | 36.06 | 997.429 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0 | 2.54
995.68 | 29.12 | 30.61 | 29.04 | 28.11 | 995.67 | 30.426 | 36.088 | 997.412 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.01 | 2.5
995.67 | 29.11 | 30.59 | 28.96 | 28.21 | 995.66 | 30.362 | 36.184 | 997.411 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.01 | 2.38
995.66 | 29.12 | 30.56 | 29.03 | 28.12 | 995.64 | 30.421 | 36.1 | 997.385 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.02 | 2.44
995.65 | 29.14 | 30.52 | 28.96 | 28.11 | 995.65 | 30.362 | 36.09 | 997.394 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0 | 2.41
995.64 | 29.12 | 30.49 | 29.01 | 28.05 | 995.64 | 30.405 | 36.031 | 997.383 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0 | 2.44
995.63 | 29.08 | 30.49 | 29.03 | 28 | 995.62 | 30.421 | 35.974 | 997.365 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0.01 | 2.49
995.61 | 29.11 | 30.44 | 28.99 | 27.97 | 995.6 | 30.389 | 35.945 | 997.348 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.01 | 2.47
995.62 | 29.12 | 30.41 | 28.88 | 28.03 | 995.59 | 30.293 | 36.006 | 997.339 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.03 | 2.38
995.6 | 29.03 | 30.48 | 28.9 | 28 | 995.59 | 30.314 | 35.981 | 997.339 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.01 | 2.48
995.6 | 28.99 | 30.54 | 28.81 | 28.15 | 995.58 | 30.239 | 36.132 | 997.326 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | 0.02 | 2.39
995.59 | 29 | 30.54 | 28.85 | 28.07 | 995.59 | 30.271 | 36.044 | 997.334 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0 | 2.47
995.6 | 29.03 | 30.45 | 28.93 | 27.99 | 995.58 | 30.335 | 35.97 | 997.327 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.02 | 2.46
995.62 | 29.05 | 30.48 | 28.89 | 27.93 | 995.6 | 30.303 | 35.905 | 997.345 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.02 | 2.55
995.6 | 29.06 | 30.44 | 28.81 | 28 | 995.59 | 30.234 | 35.979 | 997.338 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.01 | 2.44
995.6 | 29.06 | 30.45 | 28.86 | 27.98 | 995.58 | 30.277 | 35.96 | 997.323 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | 0.02 | 2.47
995.59 | 29.13 | 30.53 | 28.94 | 27.84 | 995.58 | 30.346 | 35.821 | 997.33 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.01 | 2.69
995.57 | 29.19 | 30.84 | 29.06 | 27.82 | 995.55 | 30.447 | 35.794 | 997.297 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.02 | 3.02
995.58 | 29.43 | 30.91 | 29.2 | 27.87 | 995.56 | 30.565 | 35.852 | 997.304 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.02 | 3.04
995.58 | 29.61 | 31.17 | 29.31 | 28.28 | 995.57 | 30.656 | 36.258 | 997.317 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0.01 | 2.89
995.57 | 29.9 | 31.2 | 29.46 | 28.39 | 995.58 | 30.779 | 36.365 | 997.324 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.44 | -0.01 | 2.81
995.57 | 30.25 | 31.26 | 29.66 | 28.52 | 995.54 | 30.949 | 36.5 | 997.282 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.59 | 0.03 | 2.74
995.58 | 30.59 | 31.19 | 29.92 | 28.81 | 995.55 | 31.168 | 36.792 | 997.299 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.67 | 0.03 | 2.38
995.57 | 30.88 | 30.98 | 30.19 | 28.96 | 995.55 | 31.398 | 36.933 | 997.301 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.69 | 0.02 | 2.02
995.56 | 31.19 | 30.46 | 30.46 | 28.64 | 995.53 | 31.622 | 36.613 | 997.272 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.73 | 0.03 | 1.82
995.55 | 31.48 | 30.2 | 30.7 | 28.57 | 995.51 | 31.825 | 36.548 | 997.261 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.78 | 0.04 | 1.63
995.53 | 31.66 | 29.71 | 30.94 | 28.42 | 995.51 | 32.028 | 36.397 | 997.256 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.72 | 0.02 | 1.29
995.52 | 31.64 | 29.33 | 31.03 | 28.15 | 995.5 | 32.098 | 36.128 | 997.246 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.61 | 0.02 | 1.18
995.52 | 31.49 | 29.23 | 30.95 | 28.11 | 995.48 | 32.034 | 36.092 | 997.23 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.54 | 0.04 | 1.12
995.51 | 31.39 | 29.11 | 30.89 | 27.99 | 995.47 | 31.98 | 35.964 | 997.219 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.5 | 0.04 | 1.12
995.5 | 31.27 | 29.02 | 30.75 | 27.95 | 995.47 | 31.863 | 35.932 | 997.219 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.52 | 0.03 | 1.07
995.5 | 31.06 | 28.87 | 30.67 | 27.65 | 995.46 | 31.799 | 35.627 | 997.204 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.39 | 0.04 | 1.22
995.48 | 30.84 | 29.04 | 30.56 | 27.46 | 995.45 | 31.708 | 35.436 | 997.195 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.03 | 1.58
995.47 | 30.59 | 29.37 | 30.46 | 27.48 | 995.44 | 31.622 | 35.461 | 997.189 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.03 | 1.89
995.48 | 30.4 | 29.63 | 30.28 | 27.74 | 995.43 | 31.473 | 35.72 | 997.174 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.05 | 1.89
995.45 | 30.14 | 29.93 | 30.03 | 28.03 | 995.46 | 31.265 | 36.002 | 997.207 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | -0.01 | 1.9
995.45 | 30 | 30.19 | 29.91 | 28.21 | 995.45 | 31.158 | 36.191 | 997.193 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0 | 1.98
995.44 | 29.83 | 30.39 | 29.82 | 28.23 | 995.45 | 31.083 | 36.203 | 997.194 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.01 | -0.01 | 2.16
995.42 | 29.76 | 30.45 | 29.75 | 28.24 | 995.45 | 31.024 | 36.222 | 997.198 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.01 | -0.03 | 2.21
995.42 | 29.67 | 30.57 | 29.67 | 28.21 | 995.43 | 30.955 | 36.191 | 997.177 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0 | -0.01 | 2.36
995.43 | 29.71 | 30.47 | 29.63 | 28.25 | 995.42 | 30.923 | 36.227 | 997.167 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.01 | 2.22
995.41 | 29.65 | 30.51 | 29.65 | 28.16 | 995.42 | 30.939 | 36.138 | 997.168 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0 | -0.01 | 2.35
995.41 | 29.61 | 30.55 | 29.54 | 28.19 | 995.39 | 30.848 | 36.166 | 997.137 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.07 | 0.02 | 2.36
995.45 | 29.6 | 30.81 | 29.52 | 28.31 | 995.39 | 30.832 | 36.29 | 997.134 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.06 | 2.5
995.43 | 29.66 | 30.74 | 29.62 | 28.33 | 995.43 | 30.912 | 36.304 | 997.172 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | 0 | 2.41
995.43 | 29.67 | 30.76 | 29.68 | 28.31 | 995.4 | 30.966 | 36.283 | 997.151 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.01 | 0.03 | 2.45
995.42 | 29.7 | 30.66 | 29.6 | 28.34 | 995.42 | 30.901 | 36.319 | 997.164 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0 | 2.32
995.42 | 29.6 | 30.82 | 29.54 | 28.42 | 995.39 | 30.848 | 36.397 | 997.138 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0.03 | 2.4
995.42 | 29.65 | 30.76 | 29.49 | 28.42 | 995.4 | 30.805 | 36.393 | 997.146 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.02 | 2.34
995.4 | 29.59 | 30.82 | 29.54 | 28.43 | 995.4 | 30.848 | 36.411 | 997.148 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0 | 2.39
995.4 | 29.58 | 30.8 | 29.52 | 28.35 | 995.39 | 30.832 | 36.325 | 997.137 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0.01 | 2.45
995.4 | 29.6 | 30.78 | 29.58 | 28.31 | 995.37 | 30.88 | 36.288 | 997.121 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.02 | 0.03 | 2.47
995.41 | 29.52 | 30.86 | 29.57 | 28.33 | 995.37 | 30.875 | 36.304 | 997.114 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.05 | 0.04 | 2.53
995.41 | 29.52 | 30.92 | 29.41 | 28.43 | 995.39 | 30.736 | 36.405 | 997.14 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.11 | 0.02 | 2.49
995.42 | 29.46 | 31.03 | 29.42 | 28.48 | 995.39 | 30.752 | 36.458 | 997.141 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | 0.03 | 2.55
995.42 | 29.73 | 30.93 | 29.47 | 28.47 | 995.39 | 30.789 | 36.449 | 997.14 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0.03 | 2.46
995.44 | 29.79 | 30.76 | 29.65 | 28.32 | 995.38 | 30.944 | 36.3 | 997.126 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.06 | 2.44
995.45 | 29.81 | 30.67 | 29.62 | 28.37 | 995.43 | 30.912 | 36.348 | 997.172 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.02 | 2.3
995.43 | 29.71 | 30.72 | 29.61 | 28.4 | 995.41 | 30.907 | 36.38 | 997.152 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.02 | 2.32
995.43 | 29.64 | 30.8 | 29.47 | 28.55 | 995.41 | 30.789 | 36.527 | 997.154 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | 0.02 | 2.25
995.41 | 29.58 | 30.86 | 29.46 | 28.51 | 995.41 | 30.784 | 36.489 | 997.156 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0 | 2.35
995.42 | 29.69 | 30.73 | 29.54 | 28.44 | 995.39 | 30.848 | 36.422 | 997.132 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.03 | 2.29
995.42 | 29.58 | 30.8 | 29.63 | 28.26 | 995.38 | 30.923 | 36.239 | 997.13 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.05 | 0.04 | 2.54
995.44 | 29.57 | 30.86 | 29.45 | 28.48 | 995.4 | 30.773 | 36.46 | 997.151 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.04 | 2.38
995.44 | 29.6 | 30.76 | 29.63 | 28.24 | 995.43 | 30.923 | 36.216 | 997.173 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.03 | 0.01 | 2.52
995.46 | 29.61 | 30.77 | 29.6 | 28.34 | 995.4 | 30.896 | 36.313 | 997.15 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.01 | 0.06 | 2.43
995.45 | 29.64 | 30.71 | 29.49 | 28.35 | 995.42 | 30.811 | 36.323 | 997.171 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | 0.03 | 2.36
995.43 | 29.73 | 30.52 | 29.6 | 28.21 | 995.44 | 30.896 | 36.185 | 997.186 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | -0.01 | 2.31
995.43 | 29.7 | 30.54 | 29.62 | 28.15 | 995.43 | 30.912 | 36.123 | 997.179 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0 | 2.39
995.44 | 29.72 | 30.53 | 29.72 | 27.99 | 995.42 | 30.998 | 35.966 | 997.169 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0 | 0.02 | 2.54
995.45 | 29.72 | 30.45 | 29.59 | 28.16 | 995.43 | 30.891 | 36.138 | 997.176 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.02 | 2.29
995.45 | 29.7 | 30.44 | 29.6 | 28.17 | 995.42 | 30.901 | 36.151 | 997.168 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.03 | 2.27
995.46 | 29.65 | 30.51 | 29.6 | 28.1 | 995.42 | 30.896 | 36.079 | 997.169 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0.04 | 2.41
995.45 | 29.57 | 30.65 | 29.53 | 28.15 | 995.44 | 30.843 | 36.128 | 997.186 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | 0.01 | 2.5
995.44 | 29.56 | 30.65 | 29.46 | 28.18 | 995.44 | 30.784 | 36.153 | 997.182 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0 | 2.47
995.46 | 29.6 | 30.58 | 29.52 | 28.18 | 995.42 | 30.832 | 36.161 | 997.163 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.04 | 2.4
995.47 | 29.54 | 30.65 | 29.44 | 28.18 | 995.45 | 30.763 | 36.161 | 997.2 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.02 | 2.47
995.46 | 29.55 | 30.64 | 29.49 | 28.11 | 995.45 | 30.811 | 36.082 | 997.192 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.06 | 0.01 | 2.53
995.49 | 29.59 | 30.57 | 29.51 | 28.04 | 995.47 | 30.827 | 36.016 | 997.215 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.02 | 2.53
995.53 | 29.56 | 30.57 | 29.44 | 28.13 | 995.47 | 30.768 | 36.103 | 997.22 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0.06 | 2.44
995.49 | 29.49 | 30.67 | 29.41 | 28.2 | 995.48 | 30.741 | 36.176 | 997.231 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.01 | 2.47
995.48 | 29.46 | 30.7 | 29.37 | 28.23 | 995.46 | 30.709 | 36.208 | 997.209 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.02 | 2.47
995.51 | 29.51 | 30.63 | 29.43 | 28.15 | 995.46 | 30.757 | 36.132 | 997.209 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | 0.05 | 2.48
995.48 | 29.49 | 30.68 | 29.39 | 28.23 | 995.49 | 30.72 | 36.21 | 997.24 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | -0.01 | 2.45
995.5 | 29.5 | 30.65 | 29.27 | 28.29 | 995.47 | 30.618 | 36.269 | 997.217 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.03 | 2.36
995.5 | 29.47 | 30.68 | 29.34 | 28.25 | 995.49 | 30.677 | 36.227 | 997.238 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.01 | 2.43
995.49 | 29.48 | 30.65 | 29.29 | 28.29 | 995.48 | 30.64 | 36.271 | 997.227 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.01 | 2.36
995.49 | 29.51 | 30.6 | 29.35 | 28.2 | 995.46 | 30.688 | 36.178 | 997.211 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.03 | 2.4
995.47 | 29.48 | 30.63 | 29.39 | 28.16 | 995.46 | 30.72 | 36.136 | 997.21 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.01 | 2.47
995.52 | 29.44 | 30.67 | 29.25 | 28.3 | 995.48 | 30.608 | 36.277 | 997.227 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | 0.04 | 2.37
995.49 | 29.45 | 30.65 | 29.3 | 28.27 | 995.5 | 30.65 | 36.245 | 997.245 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | -0.01 | 2.38
995.5 | 29.47 | 30.6 | 29.37 | 28.05 | 995.49 | 30.704 | 36.025 | 997.232 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.1 | 0.01 | 2.55
995.53 | 29.44 | 30.63 | 29.23 | 28.32 | 995.5 | 30.586 | 36.296 | 997.248 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.03 | 2.31
995.52 | 29.37 | 30.74 | 29.14 | 28.4 | 995.52 | 30.512 | 36.38 | 997.269 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0 | 2.34
995.51 | 29.41 | 30.76 | 29.25 | 28.26 | 995.51 | 30.602 | 36.233 | 997.252 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0 | 2.5
995.52 | 29.52 | 30.65 | 29.24 | 28.27 | 995.5 | 30.597 | 36.25 | 997.251 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.02 | 2.38
995.54 | 29.56 | 30.61 | 29.35 | 28.22 | 995.53 | 30.688 | 36.195 | 997.279 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.01 | 2.39
995.55 | 29.6 | 30.54 | 29.21 | 28.44 | 995.53 | 30.57 | 36.412 | 997.273 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.39 | 0.02 | 2.1
995.56 | 29.62 | 30.51 | 29.27 | 28.32 | 995.54 | 30.624 | 36.296 | 997.29 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.35 | 0.02 | 2.19
995.55 | 29.61 | 30.47 | 29.49 | 28.01 | 995.55 | 30.811 | 35.991 | 997.295 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.12 | 0 | 2.46
995.56 | 29.53 | 30.65 | 29.44 | 28.1 | 995.55 | 30.763 | 36.081 | 997.297 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.01 | 2.55
995.57 | 29.58 | 30.74 | 29.37 | 28.28 | 995.56 | 30.704 | 36.262 | 997.304 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.01 | 2.46
995.56 | 29.74 | 30.71 | 29.41 | 28.4 | 995.56 | 30.736 | 36.382 | 997.311 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.33 | 0 | 2.31
995.58 | 29.75 | 31.16 | 29.46 | 28.51 | 995.53 | 30.784 | 36.483 | 997.278 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.29 | 0.05 | 2.65
995.59 | 29.79 | 31.25 | 29.49 | 28.89 | 995.57 | 30.805 | 36.868 | 997.319 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0.02 | 2.36
995.58 | 29.81 | 31.54 | 29.45 | 29.04 | 995.57 | 30.773 | 37.015 | 997.32 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.36 | 0.01 | 2.5
995.57 | 29.84 | 31.87 | 29.57 | 29.51 | 995.58 | 30.875 | 37.486 | 997.33 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | -0.01 | 2.36
995.56 | 29.9 | 31.92 | 29.54 | 29.85 | 995.57 | 30.851 | 37.824 | 997.319 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.36 | -0.01 | 2.07
995.55 | 29.91 | 31.96 | 29.66 | 29.52 | 995.58 | 30.952 | 37.494 | 997.327 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | -0.03 | 2.44
995.56 | 29.87 | 32.37 | 29.62 | 30.08 | 995.55 | 30.915 | 38.053 | 997.3 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | 0.01 | 2.29
995.53 | 30.14 | 32.64 | 29.88 | 30.7 | 995.58 | 31.134 | 38.675 | 997.326 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | -0.05 | 1.94
995.53 | 30.28 | 32.63 | 30.11 | 30.25 | 995.55 | 31.331 | 38.228 | 997.292 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.17 | -0.02 | 2.38
995.53 | 30.37 | 32.31 | 30.16 | 30.32 | 995.55 | 31.369 | 38.293 | 997.292 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | -0.02 | 1.99
995.52 | 30.44 | 32.74 | 30.26 | 29.89 | 995.53 | 31.454 | 37.87 | 997.279 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.18 | -0.01 | 2.85
995.52 | 30.5 | 32.58 | 30.41 | 30.05 | 995.51 | 31.582 | 38.032 | 997.255 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.01 | 2.53
995.52 | 30.44 | 32.56 | 30.39 | 30.18 | 995.49 | 31.566 | 38.162 | 997.232 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0.03 | 2.38
995.52 | 30.27 | 33.02 | 30.32 | 30.62 | 995.49 | 31.508 | 38.593 | 997.24 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.05 | 0.03 | 2.4
995.51 | 30.19 | 33.16 | 30.22 | 30.69 | 995.5 | 31.422 | 38.669 | 997.248 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.03 | 0.01 | 2.47
995.5 | 30.15 | 33.16 | 30.23 | 30.59 | 995.48 | 31.427 | 38.564 | 997.229 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.08 | 0.02 | 2.57
995.5 | 30.07 | 33.41 | 30.13 | 30.75 | 995.49 | 31.342 | 38.732 | 997.24 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.06 | 0.01 | 2.66
995.46 | 30.14 | 33.55 | 30.15 | 30.71 | 995.48 | 31.358 | 38.692 | 997.227 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | -0.01 | -0.02 | 2.84
995.46 | 30.11 | 33.28 | 30.06 | 30.95 | 995.45 | 31.289 | 38.924 | 997.197 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.05 | 0.01 | 2.33
995.45 | 30.04 | 33.57 | 30 | 30.63 | 995.48 | 31.235 | 38.608 | 997.226 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | -0.03 | 2.94
995.43 | 30.11 | 33.91 | 30.02 | 30.94 | 995.46 | 31.251 | 38.917 | 997.203 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | -0.03 | 2.97
995.42 | 30.12 | 33.55 | 30.08 | 31.76 | 995.44 | 31.299 | 39.733 | 997.191 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | -0.02 | 1.79
995.41 | 30.07 | 33.76 | 30.06 | 31.32 | 995.42 | 31.289 | 39.298 | 997.167 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.01 | -0.01 | 2.44
995.4 | 30.03 | 33.86 | 29.99 | 31.09 | 995.42 | 31.23 | 39.069 | 997.168 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.04 | -0.02 | 2.77
995.36 | 30.09 | 33.6 | 30 | 31.23 | 995.4 | 31.235 | 39.207 | 997.142 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | -0.04 | 2.37
995.38 | 30.05 | 33.74 | 29.92 | 31.41 | 995.37 | 31.166 | 39.388 | 997.117 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | 0.01 | 2.33
995.36 | 30.08 | 33.78 | 29.92 | 31.3 | 995.39 | 31.171 | 39.279 | 997.141 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | -0.03 | 2.48
995.37 | 30.04 | 33.88 | 29.95 | 31.26 | 995.36 | 31.192 | 39.237 | 997.11 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.09 | 0.01 | 2.62
995.38 | 30.05 | 33.7 | 29.91 | 31.58 | 995.36 | 31.16 | 39.562 | 997.109 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | 0.02 | 2.12
995.38 | 30.01 | 33.77 | 29.87 | 31.43 | 995.39 | 31.128 | 39.407 | 997.139 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | -0.01 | 2.34
995.38 | 30.03 | 34.08 | 29.84 | 31.71 | 995.4 | 31.102 | 39.687 | 997.145 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | -0.02 | 2.37
995.37 | 29.96 | 33.86 | 29.88 | 31.39 | 995.4 | 31.134 | 39.363 | 997.143 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.08 | -0.03 | 2.47
995.36 | 30 | 33.98 | 29.78 | 31.92 | 995.38 | 31.048 | 39.899 | 997.13 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | -0.02 | 2.06
995.35 | 29.97 | 34.17 | 29.77 | 32.08 | 995.38 | 31.043 | 40.057 | 997.129 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | -0.03 | 2.09
995.34 | 29.95 | 34.05 | 29.74 | 31.7 | 995.37 | 31.016 | 39.674 | 997.115 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | -0.03 | 2.35
995.33 | 29.9 | 34.09 | 29.62 | 31.88 | 995.35 | 30.915 | 39.853 | 997.097 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | -0.02 | 2.21
995.37 | 29.9 | 34.18 | 29.74 | 31.64 | 995.35 | 31.016 | 39.621 | 997.097 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.16 | 0.02 | 2.54
995.38 | 29.96 | 34.09 | 29.73 | 31.57 | 995.39 | 31.011 | 39.546 | 997.135 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | -0.01 | 2.52
995.37 | 29.95 | 34.05 | 29.69 | 31.61 | 995.39 | 30.979 | 39.584 | 997.14 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | -0.02 | 2.44
995.37 | 29.94 | 34.18 | 29.75 | 31.6 | 995.38 | 31.027 | 39.573 | 997.122 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | -0.01 | 2.58
995.37 | 29.93 | 34.3 | 29.66 | 32.11 | 995.38 | 30.952 | 40.082 | 997.129 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | -0.01 | 2.19
995.38 | 29.94 | 34.38 | 29.72 | 32.09 | 995.38 | 31 | 40.063 | 997.126 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | 0 | 2.29
995.39 | 29.95 | 34.5 | 29.65 | 31.97 | 995.39 | 30.941 | 39.943 | 997.139 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0 | 2.53
995.4 | 29.93 | 34.63 | 29.78 | 32.02 | 995.41 | 31.054 | 39.998 | 997.152 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.15 | -0.01 | 2.61
995.38 | 29.92 | 34.65 | 29.78 | 32.3 | 995.43 | 31.054 | 40.277 | 997.174 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | -0.05 | 2.35
995.4 | 29.91 | 34.95 | 29.66 | 32.32 | 995.42 | 30.947 | 40.302 | 997.162 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | -0.02 | 2.63
995.39 | 29.9 | 34.83 | 29.7 | 32.39 | 995.42 | 30.984 | 40.365 | 997.167 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.2 | -0.03 | 2.44
995.38 | 29.89 | 35.04 | 29.75 | 32.27 | 995.39 | 31.022 | 40.25 | 997.139 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | -0.01 | 2.77
995.35 | 29.97 | 34.23 | 29.69 | 32.5 | 995.39 | 30.974 | 40.477 | 997.138 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | -0.04 | 1.73
995.37 | 29.89 | 34.12 | 29.75 | 31.76 | 995.38 | 31.022 | 39.739 | 997.128 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.14 | -0.01 | 2.36
995.38 | 29.85 | 34.1 | 29.61 | 31.97 | 995.39 | 30.904 | 39.943 | 997.132 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | -0.01 | 2.13
995.38 | 29.89 | 33.98 | 29.56 | 31.83 | 995.39 | 30.867 | 39.809 | 997.135 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.33 | -0.01 | 2.15
995.4 | 29.87 | 33.85 | 29.64 | 31.65 | 995.4 | 30.931 | 39.623 | 997.143 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0 | 2.2
995.37 | 29.91 | 33.67 | 29.62 | 31.55 | 995.4 | 30.92 | 39.531 | 997.146 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.29 | -0.03 | 2.12
995.39 | 29.88 | 33.69 | 29.57 | 31.53 | 995.39 | 30.877 | 39.504 | 997.133 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.31 | 0 | 2.16
995.36 | 29.86 | 33.57 | 29.59 | 31.46 | 995.39 | 30.888 | 39.441 | 997.139 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | -0.03 | 2.11
995.36 | 29.83 | 33.55 | 29.55 | 31.39 | 995.39 | 30.861 | 39.369 | 997.132 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | -0.03 | 2.16
995.35 | 29.84 | 33.45 | 29.51 | 31.39 | 995.35 | 30.824 | 39.363 | 997.101 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.33 | 0 | 2.06
995.38 | 29.85 | 33.35 | 29.51 | 31.24 | 995.35 | 30.824 | 39.218 | 997.101 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.34 | 0.03 | 2.11
995.36 | 29.89 | 33.28 | 29.48 | 31.21 | 995.38 | 30.797 | 39.186 | 997.128 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.41 | -0.02 | 2.07
995.38 | 29.91 | 33.14 | 29.61 | 30.94 | 995.38 | 30.909 | 38.921 | 997.124 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0 | 2.2
995.35 | 29.87 | 33.1 | 29.54 | 31.13 | 995.37 | 30.845 | 39.104 | 997.113 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.33 | -0.02 | 1.97
995.34 | 29.88 | 33.06 | 29.59 | 31.03 | 995.34 | 30.888 | 39.003 | 997.088 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.29 | 0 | 2.03
995.35 | 29.83 | 33.02 | 29.57 | 30.86 | 995.33 | 30.877 | 38.837 | 997.077 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0.02 | 2.16
995.36 | 29.83 | 33.02 | 29.6 | 30.81 | 995.35 | 30.899 | 38.787 | 997.099 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.01 | 2.21
995.35 | 29.74 | 33.03 | 29.48 | 30.85 | 995.35 | 30.803 | 38.825 | 997.095 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.26 | 0 | 2.18
995.33 | 29.75 | 33.05 | 29.42 | 30.95 | 995.35 | 30.744 | 38.923 | 997.099 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.33 | -0.02 | 2.1
995.32 | 29.75 | 32.99 | 29.52 | 30.72 | 995.34 | 30.835 | 38.699 | 997.085 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | -0.02 | 2.27
995.33 | 29.75 | 32.94 | 29.47 | 30.77 | 995.32 | 30.787 | 38.743 | 997.064 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.28 | 0.01 | 2.17
995.34 | 29.68 | 32.95 | 29.45 | 30.66 | 995.33 | 30.771 | 38.64 | 997.074 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.23 | 0.01 | 2.29
995.28 | 29.7 | 32.93 | 29.4 | 30.72 | 995.35 | 30.733 | 38.701 | 997.095 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | -0.07 | 2.21
995.29 | 29.68 | 32.89 | 29.41 | 30.73 | 995.28 | 30.739 | 38.707 | 997.025 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | 0.01 | 2.16
995.27 | 29.63 | 32.9 | 29.41 | 30.61 | 995.28 | 30.739 | 38.583 | 997.027 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.22 | -0.01 | 2.29
995.26 | 29.62 | 32.88 | 29.41 | 30.57 | 995.27 | 30.739 | 38.543 | 997.012 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | -0.01 | 2.31
995.24 | 29.6 | 32.84 | 29.47 | 30.48 | 995.25 | 30.787 | 38.455 | 996.999 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.13 | -0.01 | 2.36
995.24 | 29.65 | 32.8 | 29.35 | 30.58 | 995.23 | 30.69 | 38.558 | 996.978 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | 0.01 | 2.22
995.25 | 29.64 | 32.73 | 29.43 | 30.47 | 995.24 | 30.76 | 38.446 | 996.989 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.21 | 0.01 | 2.26
995.22 | 29.61 | 32.76 | 29.42 | 30.42 | 995.26 | 30.749 | 38.396 | 997.011 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.19 | -0.04 | 2.34
995.21 | 29.6 | 32.72 | 29.36 | 30.44 | 995.23 | 30.701 | 38.417 | 996.98 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | -0.02 | 2.28
995.18 | 29.57 | 32.76 | 29.32 | 30.38 | 995.22 | 30.664 | 38.36 | 996.971 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.25 | -0.04 | 2.38
995.23 | 29.54 | 32.75 | 29.3 | 30.36 | 995.19 | 30.648 | 38.335 | 996.939 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | 0.04 | 2.39
995.21 | 29.6 | 32.67 | 29.31 | 30.46 | 995.23 | 30.653 | 38.442 | 996.973 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.29 | -0.02 | 2.21
995.16 | 29.61 | 32.61 | 29.32 | 30.33 | 995.2 | 30.664 | 38.303 | 996.942 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.29 | -0.04 | 2.28
995.16 | 29.56 | 32.61 | 29.26 | 30.37 | 995.18 | 30.616 | 38.35 | 996.927 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | -0.02 | 2.24
995.16 | 29.55 | 32.59 | 29.31 | 30.27 | 995.17 | 30.658 | 38.244 | 996.914 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.24 | -0.01 | 2.32
995.13 | 29.61 | 32.56 | 29.22 | 30.44 | 995.16 | 30.578 | 38.417 | 996.904 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.39 | -0.03 | 2.12
995.13 | 29.61 | 32.51 | 29.27 | 30.31 | 995.13 | 30.621 | 38.287 | 996.874 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.34 | 0 | 2.2
995.1 | 29.62 | 32.47 | 29.22 | 30.31 | 995.15 | 30.578 | 38.289 | 996.894 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.4 | -0.05 | 2.16
995.04 | 29.54 | 32.5 | 29.24 | 30.31 | 995.1 | 30.6 | 38.285 | 996.848 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.3 | -0.06 | 2.19
995.09 | 29.59 | 32.42 | 29.19 | 30.39 | 995.04 | 30.557 | 38.364 | 996.788 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.4 | 0.05 | 2.03
995.12 | 29.54 | 32.46 | 29.25 | 30.24 | 995.09 | 30.605 | 38.215 | 996.84 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.29 | 0.03 | 2.22
995.11 | 29.5 | 32.5 | 29.21 | 30.26 | 995.09 | 30.568 | 38.234 | 996.841 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.29 | 0.02 | 2.24
995.1 | 29.49 | 32.52 | 29.22 | 30.2 | 995.1 | 30.584 | 38.179 | 996.851 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.27 | 0 | 2.32
995.08 | 29.48 | 32.52 | 29.14 | 30.29 | 995.1 | 30.514 | 38.265 | 996.85 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.34 | -0.02 | 2.23
995.07 | 29.48 | 32.46 | 29.17 | 30.2 | 995.09 | 30.541 | 38.173 | 996.834 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.31 | -0.02 | 2.26
995.05 | 29.46 | 32.44 | 29.15 | 30.25 | 995.06 | 30.52 | 38.232 | 996.811 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.31 | -0.01 | 2.19
995.02 | 29.35 | 32.64 | 28.98 | 30.33 | 995.06 | 30.375 | 38.303 | 996.808 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.37 | -0.04 | 2.31
995.04 | 29.33 | 32.66 | 28.95 | 30.45 | 995.04 | 30.349 | 38.423 | 996.783 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.38 | 0 | 2.21
995.03 | 29.43 | 32.63 | 28.99 | 30.42 | 995.05 | 30.386 | 38.4 | 996.794 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.44 | -0.02 | 2.21
995.01 | 29.53 | 32.5 | 29.14 | 30.21 | 995.01 | 30.509 | 38.182 | 996.753 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.39 | 0 | 2.29
994.99 | 29.62 | 32.38 | 29.14 | 30.22 | 994.99 | 30.509 | 38.2 | 996.739 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.48 | 0 | 2.16
994.99 | 29.54 | 32.44 | 29.15 | 30.22 | 994.98 | 30.525 | 38.202 | 996.728 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.39 | 0.01 | 2.22
994.97 | 29.44 | 32.6 | 29.02 | 30.38 | 994.98 | 30.413 | 38.354 | 996.726 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.42 | -0.01 | 2.22
994.97 | 29.42 | 32.62 | 29.06 | 30.36 | 994.98 | 30.445 | 38.333 | 996.722 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.36 | -0.01 | 2.26
994.98 | 29.4 | 32.71 | 28.93 | 30.57 | 995 | 30.338 | 38.545 | 996.746 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.47 | -0.02 | 2.14
994.94 | 29.47 | 32.63 | 29.05 | 30.36 | 994.98 | 30.434 | 38.337 | 996.726 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.42 | -0.04 | 2.27
994.95 | 29.44 | 32.69 | 29.01 | 30.53 | 994.94 | 30.402 | 38.505 | 996.685 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.43 | 0.01 | 2.16
994.94 | 29.39 | 32.75 | 29.05 | 30.47 | 994.94 | 30.439 | 38.451 | 996.691 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.34 | 0 | 2.28
994.92 | 29.4 | 32.75 | 29.03 | 30.45 | 994.96 | 30.418 | 38.423 | 996.704 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.37 | -0.04 | 2.3
994.93 | 29.4 | 32.76 | 29.02 | 30.44 | 994.95 | 30.413 | 38.419 | 996.695 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.38 | -0.02 | 2.32
994.92 | 29.4 | 32.75 | 29.05 | 30.47 | 994.95 | 30.434 | 38.446 | 996.696 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.35 | -0.03 | 2.28
994.89 | 29.38 | 32.76 | 29.05 | 30.51 | 994.92 | 30.439 | 38.484 | 996.669 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.33 | -0.03 | 2.25
994.9 | 29.32 | 32.89 | 28.93 | 30.71 | 994.91 | 30.333 | 38.684 | 996.653 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.39 | -0.01 | 2.18
994.9 | 29.37 | 32.83 | 28.89 | 30.72 | 994.9 | 30.306 | 38.696 | 996.647 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.48 | 0 | 2.11
994.85 | 29.33 | 32.86 | 28.96 | 30.69 | 994.91 | 30.359 | 38.663 | 996.656 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.37 | -0.06 | 2.17
994.87 | 29.31 | 32.92 | 28.88 | 30.66 | 994.87 | 30.29 | 38.64 | 996.612 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.43 | 0 | 2.26
994.82 | 29.28 | 32.94 | 28.89 | 30.72 | 994.88 | 30.306 | 38.699 | 996.631 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.39 | -0.06 | 2.22
994.83 | 29.32 | 32.95 | 28.86 | 30.73 | 994.84 | 30.279 | 38.703 | 996.587 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.46 | -0.01 | 2.22
994.85 | 29.35 | 32.9 | 28.83 | 30.81 | 994.85 | 30.253 | 38.787 | 996.594 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.52 | 0 | 2.09
994.86 | 29.36 | 32.86 | 28.82 | 30.8 | 994.87 | 30.242 | 38.781 | 996.613 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.54 | -0.01 | 2.06
994.84 | 29.33 | 32.89 | 28.96 | 30.6 | 994.86 | 30.359 | 38.577 | 996.602 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.37 | -0.02 | 2.29
994.82 | 29.34 | 32.89 | 28.91 | 30.72 | 994.86 | 30.317 | 38.696 | 996.602 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.43 | -0.04 | 2.17
994.82 | 29.39 | 32.8 | 28.88 | 30.79 | 994.85 | 30.29 | 38.762 | 996.599 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.51 | -0.03 | 2.01
994.84 | 29.32 | 32.87 | 28.97 | 30.49 | 994.81 | 30.37 | 38.469 | 996.558 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.35 | 0.03 | 2.38
994.8 | 29.3 | 32.93 | 28.84 | 30.79 | 994.84 | 30.263 | 38.762 | 996.587 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.46 | -0.04 | 2.14
994.77 | 29.34 | 32.87 | 28.86 | 30.71 | 994.81 | 30.279 | 38.692 | 996.552 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.48 | -0.04 | 2.16
994.82 | 29.31 | 32.87 | 28.88 | 30.74 | 994.78 | 30.29 | 38.715 | 996.525 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.43 | 0.04 | 2.13
994.86 | 29.38 | 32.95 | 28.91 | 30.69 | 994.82 | 30.322 | 38.669 | 996.564 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.47 | 0.04 | 2.26
994.82 | 29.53 | 32.94 | 28.91 | 30.73 | 994.89 | 30.322 | 38.709 | 996.634 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.62 | -0.07 | 2.21
994.8 | 29.67 | 32.82 | 29.05 | 30.79 | 994.83 | 30.434 | 38.772 | 996.574 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.62 | -0.03 | 2.03
994.81 | 29.74 | 32.78 | 29.18 | 30.65 | 994.81 | 30.546 | 38.625 | 996.554 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.56 | 0 | 2.13
994.75 | 29.86 | 32.67 | 29.2 | 30.76 | 994.78 | 30.562 | 38.736 | 996.528 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.66 | -0.03 | 1.91
994.77 | 29.89 | 32.58 | 29.24 | 30.78 | 994.76 | 30.6 | 38.759 | 996.509 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.65 | 0.01 | 1.8
994.77 | 29.84 | 32.62 | 29.24 | 30.8 | 994.75 | 30.6 | 38.776 | 996.499 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.6 | 0.02 | 1.82
994.76 | 29.74 | 32.85 | 29.17 | 30.84 | 994.77 | 30.541 | 38.818 | 996.515 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.57 | -0.01 | 2.01
994.78 | 29.83 | 32.91 | 29.24 | 30.92 | 994.77 | 30.594 | 38.898 | 996.517 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.59 | 0.01 | 1.99
994.8 | 29.93 | 32.69 | 29.39 | 30.75 | 994.79 | 30.723 | 38.73 | 996.536 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.54 | 0.01 | 1.94
994.81 | 29.87 | 32.72 | 29.25 | 30.97 | 994.8 | 30.605 | 38.945 | 996.55 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.62 | 0.01 | 1.75
994.81 | 29.79 | 32.78 | 29.3 | 30.9 | 994.82 | 30.648 | 38.875 | 996.564 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.49 | -0.01 | 1.88
994.78 | 29.61 | 33.01 | 29.14 | 31.06 | 994.81 | 30.514 | 39.041 | 996.553 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.47 | -0.03 | 1.95
994.79 | 29.44 | 33.26 | 29.04 | 31.17 | 994.79 | 30.429 | 39.151 | 996.536 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.4 | 0 | 2.09
994.78 | 29.36 | 33.41 | 28.88 | 31.43 | 994.81 | 30.29 | 39.405 | 996.56 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.48 | -0.03 | 1.98
994.76 | 29.43 | 33.69 | 28.84 | 31.42 | 994.81 | 30.263 | 39.401 | 996.555 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.59 | -0.05 | 2.27
994.77 | 29.78 | 33.07 | 29.03 | 31.28 | 994.8 | 30.423 | 39.254 | 996.542 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.75 | -0.03 | 1.79
994.79 | 30.06 | 32.61 | 29.26 | 31.04 | 994.78 | 30.61 | 39.016 | 996.53 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.8 | 0.01 | 1.57
994.79 | 30.38 | 32.12 | 29.43 | 30.74 | 994.79 | 30.755 | 38.715 | 996.538 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.95 | 0 | 1.38
994.78 | 30.69 | 31.63 | 29.73 | 30.38 | 994.78 | 31.011 | 38.352 | 996.529 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.96 | 0 | 1.25
994.75 | 30.98 | 31.16 | 30.02 | 29.94 | 994.76 | 31.251 | 37.915 | 996.504 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.96 | -0.01 | 1.22
994.77 | 31.24 | 30.78 | 30.32 | 29.63 | 994.73 | 31.508 | 37.603 | 996.474 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.92 | 0.04 | 1.15
994.79 | 31.47 | 30.38 | 30.55 | 29.3 | 994.75 | 31.695 | 37.276 | 996.496 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.92 | 0.04 | 1.08
994.8 | 31.71 | 29.96 | 30.79 | 28.99 | 994.74 | 31.897 | 36.967 | 996.485 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.92 | 0.06 | 0.97
994.75 | 31.93 | 29.61 | 31.05 | 28.68 | 994.74 | 32.116 | 36.659 | 996.489 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.88 | 0.01 | 0.93
994.75 | 32.12 | 29.35 | 31.22 | 28.43 | 994.73 | 32.261 | 36.407 | 996.474 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.9 | 0.02 | 0.92
994.73 | 32.28 | 29.11 | 31.41 | 28.22 | 994.74 | 32.421 | 36.197 | 996.483 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.87 | -0.01 | 0.89
994.71 | 32.4 | 28.9 | 31.58 | 28.11 | 994.71 | 32.565 | 36.088 | 996.453 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.82 | 0 | 0.79
994.67 | 32.53 | 28.71 | 31.68 | 27.91 | 994.66 | 32.645 | 35.89 | 996.41 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.85 | 0.01 | 0.8
994.68 | 32.66 | 28.54 | 31.82 | 27.77 | 994.64 | 32.768 | 35.743 | 996.383 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.84 | 0.04 | 0.77
994.68 | 32.76 | 28.37 | 31.99 | 27.62 | 994.65 | 32.912 | 35.594 | 996.4 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.77 | 0.03 | 0.75
994.65 | 32.84 | 28.25 | 32.1 | 27.5 | 994.63 | 32.998 | 35.476 | 996.38 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.74 | 0.02 | 0.75
994.71 | 32.92 | 28.13 | 32.15 | 27.41 | 994.62 | 33.04 | 35.388 | 996.365 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.77 | 0.09 | 0.72
994.72 | 33 | 28.01 | 32.24 | 27.36 | 994.68 | 33.12 | 35.337 | 996.422 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.76 | 0.04 | 0.65
994.71 | 33.07 | 27.96 | 32.35 | 27.23 | 994.69 | 33.211 | 35.211 | 996.438 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.72 | 0.02 | 0.73
994.71 | 33.12 | 27.84 | 32.41 | 27.15 | 994.67 | 33.265 | 35.123 | 996.417 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.71 | 0.04 | 0.69
994.68 | 33.17 | 27.78 | 32.46 | 27.04 | 994.66 | 33.307 | 35.018 | 996.407 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.71 | 0.02 | 0.74
994.67 | 33.21 | 27.71 | 32.51 | 27.04 | 994.65 | 33.345 | 35.016 | 996.393 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.7 | 0.02 | 0.67
994.67 | 33.26 | 27.83 | 32.6 | 27.01 | 994.65 | 33.42 | 34.99 | 996.398 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.66 | 0.02 | 0.82
994.68 | 33.35 | 27.76 | 32.7 | 26.9 | 994.65 | 33.505 | 34.879 | 996.399 | 1.1893 | -7.144 | 1 | -7.977 | 1 | -1.747 | 0.65 | 0.03 | 0.86
[15 empty] | Teplota [C] | Tlak [hPa] | Vlhkosť [%]
[14 empty] | AVG | 0.19668 | 0.041802 | 0.84924
[14 empty] | RMSE | 0.40908 | 0.053703 | 1.2827
[14 empty] | STD DEV | 0.3587 | 0.033714 | 0.96128
This chart isn’t available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.